Cyclic fatigue in water of veneer–framework composites for all-ceramic dental bridges

André R. Studarta, , , Frank Filsera, Peter Kochera, Heinz Lüthyb and Ludwig J. Gaucklera 
aDepartment of Materials, ETH-Zürich, Zurich, Switzerland
bDental School, University of Basel, Basel, Switzerland 
Received 3 August 2005;  accepted 9 January 2006.  Available online 21 February 2006.

- Abstract

- Objectives

Ceramic materials applied in dentistry may exhibit significant subcritical crack growth due to the severe cyclic loading in the aqueous environment encountered in the mouth during mastication. The authors report on the subcritical crack growth behavior of three dental restoration systems (Empress 2/IPS Eris, TZP/Cercon S and Inceram-Zirconia/Vita VM7) under cyclic loading in water, in order to establish guidelines for the use and design of long-lifetime all-ceramic posterior bridges.

- Methods

Inert strength and lifetime tests under cyclic loading in an aqueous environment were performed in a mechanical bending apparatus and evaluated with Weibull statistics.

- Results

Subcritical crack growth occurred predominantly in the outer veneer layer of the veneer–framework composites. The apatite-based veneer (IPS Eris) was more susceptible to subcritical crack propagation compared to the feldspathic glass veneers (Cercon S and Vita VM7).

- Significance

Dental restoration systems containing apatite-based veneers and weak frameworks (Empress 2/IPS Eris) are not recommended for the fabrication of all-ceramic bridges in the molar region. Conversely, veneer–framework systems consisting of feldspathic glass veneers and tough zirconia-based frameworks (TZP/Cercon S and Inceram-Zirconia/Vita VM7) may exhibit lifetimes longer than 20 years if the bridge connector is properly designed.

Keywords: Dental bridge; Fixed partial denture; Ceramics; Zirconia; Alumina; Porcelain; Cyclic fatigue; Subcritical crack growth; Lifetime

- Article Outline

1. Introduction

2. Materials and methods

3. Results and discussion

4. Conclusions

Acknowledgements

Appendix A. Appendix

References

- 1. Introduction

Layered ceramic composites have been extensively used in the last few years as single crowns and multi-unit bridges in restorative dentistry. The replacement of traditional metal-based restorations with all-ceramic prostheses has been driven by the improved esthetics and lower allergenic potential achieved using ceramic materials as core frameworks [1], [2] and [3].

However, the brittle nature and typically low reliability of ceramic materials requires a thorough analysis of the mechanical and fatigue behavior of ceramic restorations in order to guarantee long-lasting bridges and crowns. Extensive studies have been carried out by Lawn et al. [4] on the mechanical and fatigue properties of dental ceramics under the typical wear and contact load conditions applied to layered crowns.

In addition to wear and contact load, multi-unit bridges are also submitted to high tensile stresses at the gingival side of the connector during mastication. Since ceramic materials are very susceptible to failure under tensile loading, all-ceramic bridges require even more stringent mechanical properties than those needed for dental crowns. Therefore, several authors have recently investigated the mechanical strength and load bearing capacity of all-ceramic bridges [5], [6], [7], [8], [9], [10], [11], [12] and [13]. According to these studies, zirconia-based ceramics are the most appropriate to withstand the high tensile stresses that occur on multi-unit bridges.

Besides the high initial mechanical strength, a key aspect that has to be considered when using ceramics as structural components is their susceptibility to subcritical crack growth phenomena. In general, subcritical crack growth in ceramic materials is caused by the water-assisted breakage of metal-oxide bonds at the crack tip under applied stress [14]. The propagating crack reduces the strength of the ceramic component, leading to failure at stress conditions at which no fracture would be expected, based solely on the initial inert strength. In the case of toughened ceramics, cyclic loading enhances the water-assisted subcritical crack propagation due to degradation of the material's toughening mechanisms [15], [16], [17] and [18]. The subcritical crack growth behavior of ceramic materials under cyclic loading can be described by the following power-law equation:

|  | (1) |
| --- | --- |

where v is the crack velocity, ΔKI the amplitude of the stress intensity factor applied (KI,max − KI,min), KI,max and KI,min the maximum and minimum stress intensity factors, R the stress ratio KI,min/KI,max, and A and n are the subcritical crack growth parameters for cyclic loading.

The aqueous environment of the mouth and cyclic loading imparted during mastication are evidently very favorable conditions for subcritical crack propagation in ceramic restorations. The evaluation of the subcritical crack growth behavior of dental ceramics under such conditions is therefore highly desirable [2] and [19]. Previous studies have focused on the properties of the veneer and framework materials separately, giving valuable information on the crack growth parameters of the individual components of the dental restoration [19], [20], [21] and [22].

However, the lifetime estimation of a dental prosthesis also requires knowledge on the subcritical crack growth behavior of the veneer and framework materials when assembled together in the final layered composite. Investigations on the veneer–framework composite render important information about the stress distribution, failure mode, crack origin, interfacial bonding and thermal stresses, which cannot be assessed by evaluating the materials separately.

A previous article on the fast fracture (KI,applied = KIC) of layered composites [23] has shown that cracks originate on, or close to, the veneer surface and propagate towards the veneer–framework interface. The crack origin is defined by the stress distribution on the composite and the flaw population introduced in the material during processing.

Under subcritical conditions (KI,applied < KIC), the initial flaw size distribution in the veneer and framework layers changes as a function of time, according to the subcritical crack growth behavior of each material. This might lead to different failure modes and crack origins than those observed in fast fracture. In dental restorations, the subcritical crack propagation in the veneer layer is expected to be enhanced by the direct contact of this material with the mouth's aqueous environment. Cyclic loading, on the other hand, is not supposed to favor crack growth in the veneer layer due to the glassy, non-toughened nature of this material [24] and [25]. In contrast, the toughened framework core is more susceptible to toughening degradation mechanisms caused by cyclic loading [15], [16], [17] and [18]. However, no significant water-assisted crack growth is expected to occur in the framework, as long as the veneer layer protects the core material from direct exposure to water. These concurrent effects have to be considered when estimating the lifetime of all-ceramic bridges, but have so far not been addressed in the literature.

The aim of this work is to investigate the lifetime and subcritical crack growth behavior of three veneer–framework dental composites that have been used or considered as potential candidates for all-ceramic dental bridges. Lifetime experiments were conducted in water and under cyclic loading, simulating the conditions encountered in the mouth during mastication. Rectangular layered samples were used to better control the layer thickness during preparation and to assess the stresses developed on the composite using simple beam theory calculations. Based on the subcritical crack growth behavior of these layered composites, guidelines can be derived for the appropriate design and selection of veneer and framework materials for all-ceramic restorations.

- 2. Materials and methods

The veneer–framework composites investigated in this study are shown in Table 1. The layered specimens (4 mm × 5 mm × 50 mm) were prepared following the multi-layering/firing steps recommended by the suppliers and described in detail in a recent publication [23].

Table 1.

Veneer–framework composites investigated in this study

|  | Inc-VM7 | Emp2-Eris | TZP-CerS |
| --- | --- | --- | --- |
| Framework material | Glass infiltrated Al2O3–ZrO2 composite (Inceram-Zirconia) | Li2O.2SiO2 glass-ceramic (Empress 2) | Yttria-stabilized ZrO2 (3Y-TZP, Cercon) |
| Veneer material | Feldspathic glass (VITA VM 7 system) | Fluorapatite glass-ceramic (IPS Eris for E2 system) | Feldspathic glass (Cercon Ceram S) |
| Supplier | VITA Zahnfabrik, Germany | Ivoclar Vivadent, Liechtenstein | Degudent, Germany |

To determine the subcritical crack growth behavior of the veneer–framework composites, approximately 30 layered samples of each composition were evaluated in terms of lifetime and inert strength, following the method described in a previous article [19].

The inert strength measurements were performed in bending mode, applying a fast loading rate (>102 MPa s−1) in order to avoid subcritical crack growth during the experiment. In this fast fracture test, the specimen is first fixed between a rigid wall and a long cantilever. A bending moment is then applied to the specimen by continuously increasing the load at the end of the cantilever.

The lifetime tests were carried out by submitting the samples to a sinusoidal cyclic load and recording the time required for specimen failure. Cyclic loading was applied in an alternating mode (ratio R = σmin/σmax = −1) at a frequency of 10 Hz, keeping a constant maximum/minimum stress on the veneer surface. In spite of the similar inert strength of the veneers, the tests were carried out at a maximum stress level of 50 MPa for the Emp2-Eris samples, and 65 MPa for the TZP-CerS and Inc-VM7 specimens. These stress levels correspond to an applied moment of approximately 700, 1030 and 1070 N mm on the Emp2-Eris, TZP-CerS and Inc-VM7 composites, respectively. The magnitude of the maximum stress was chosen based on preliminary tests, from which the stress level necessary to induce sample failure within a reasonable amount of time (<11 days) was established.

More details on the bending machine and testing configurations used for the above measurements can be found in reference [19]. The inert strength and lifetime data obtained by this means were subsequently evaluated using Weibull statistics.

- 3. Results and discussion

The fracture of the layered composites under fast loading conditions was initiated from natural cracks on the veneer layer which propagated towards the veneer–framework interface, as discussed in a preceding investigation [23]. In composites containing a tough core framework (TZP-CerS and Inc-VM7), the crack propagating from the veneer layer was deflected at the interface resembling the crack arresting mechanism that takes place at the dentin–enamel junction of natural teeth [26]. Fig. 1 shows the stress required for crack propagation from the veneer layer estimated from beam theory calculations [23]. The characteristic strength (σc,0) and Weibull modulus (m) of the veneers did not vary considerably among the investigated composites (Table 2), with values within the typically low strength level of porcelain glasses.

|  | Display Full Size version of this image (21K) |
| --- | --- |

Fig. 1. Weibull distribution for the inert strength of the veneer layer of the veneer–framework composites, according to Studart et al. [23].

Table 2.

Weibull parameters obtained for the inert strength and lifetime of the veneer material of layered specimens

|  | VM7 in Inc-VM7 | Eris in Emp2-Eris | CerS in TZP-CerS |
| --- | --- | --- | --- |
| Inert strength |  |  |  |
| m | 8.9 | 9.5 | 9.2 |
| σc,0 (MPa) | 107.4 | 96.6 | 103.9 |
| Lifetime |  |  |  |
| σmax applied (MPa) | 65 | 50 | 65 |
| m* | 0.34 | 1.48 | 0.39 |
| Nf,0 (cycles) | 490050 | 3210 | 73690 |

The lifetime experiments were performed by applying a cyclic bending load with fixed maximum stress to the veneer surface. All specimens tested under wet cyclic loading fractured from cracks originating on the veneer outer layer. The precise crack origin was not as evident as in the fast fractured samples, due to the lower stress magnitude applied in the lifetime experiments. These results indicate that subcritical crack propagation under conditions similar to those encountered in the mouth leads to the fracture of the veneer layer before the rupture of the core framework material. The water-assisted fracture of the veneer material therefore occurs faster than the cyclic load-induced rupture of the inner framework material.

Lifetime data obtained for the veneer–framework composites under cyclic loading in water are shown in Fig. 2. Despite the lower maximum stress applied (50 MPa), the Emp2-Eris composite showed a surprisingly low resistance to subcritical crack growth, exhibiting lifetimes markedly shorter than those obtained for the TZP-CerS and Inc-VM7 specimens. The fracture of the Emp2-Eris samples occurred in the same mode as that observed in the fast loading experiments [23], with the crack propagating from the veneer layer throughout the composite without any crack deflection at the veneer–framework interface. The lifetime of the Emp2-Eris composites was thus exclusively determined by the subcritical crack growth behavior of the veneer material, indicating that these restorations are highly susceptible to complete failure at rather low stress conditions (50 MPa).

|  | Display Full Size version of this image (25K) |
| --- | --- |

Fig. 2. Weibull distribution for the number of cycles to failure of the veneer layer of the veneer–framework composites.

In the zirconia-based composites (TZP-CerS and Inc-VM7), on the other hand, cracks were deflected at the veneer–framework interface as in the fast fracture tests [23]. Due to crack deflection at the interface, the initial cracks in the veneer layer of the Inc-VM7 and TZP-CerS composites do not lead to complete failure of the all-ceramic restoration, but only exposes the framework core to the aqueous environment. The lifetime data shown in Fig. 2 refer only to the fracture of the veneer material, since the stress level applied was not high enough to break the zirconia-based framework core after crack deflection at the interface. The lifetime of the core framework under the aqueous and cyclic conditions expected after the failure of the veneer layer will be addressed in a following publication.

The veneer layer of the TZP-CerS and Inc-VM7 composites exhibited lifetimes markedly longer than that obtained for the Emp2-Eris specimens, in spite of the higher maximum stress applied during the test (65 MPa). In the case of the Inc-VM7 composite, the veneer layer of some of the specimens did not crack even within 11 days of testing (Fig. 2). Since the probability of failure due to subcritical crack growth is determined by the initial flaw population in the material (in this case the veneer), the lifetime data can also be described in terms of Weibull statistics [19], as shown in Fig. 2 andTable 2.

The Weibull lifetime data depicted in Fig. 2 and the inert strength results shown in Fig. 1 can be combined to derive the subcritical crack growth parameters A and n of the veneer materials (Eq. (1)), using the following equations [19] and [25]:

|  | (2) |
| --- | --- |

and

|  | (3) |
| --- | --- |

where m* and Nf,0 are the Weibull modulus and the characteristic number of cycles to failure from the lifetime distribution (Fig. 2 and Table 2), KIC the critical stress intensity factor, Y the geometrical constant related to the crack location and shape (equal to 1.3 for surface cracks) and Δσ is the amplitude of the applied stress. Critical stress intensity factors (KIC) reported in the literature for apatite (1.0 MPa m1/2) [27] and for feldspathic porcelains (0.9 MPa m1/2) [28] were used to calculate the parameters A and n for the Eris, and the VM7 and CerS veneers, respectively.

It is important to note that the parameters A and n can only be calculated from Eqs. (2) and (3) when the fractures obtained in the inert strength and the lifetime measurements are caused by the same flaw population within the material [18], [25], [29] and [30]. Even though it was not possible to identify the precise origin of cracks in specimens submitted to the lifetime tests, it is reasonable to assume that in this case the initial cracks propagated from the veneer surface where the stress is highest and the material is directly exposed to water. For the strength measurements, on the other hand, only part (40–60%) of the specimens fractured from flaws on the veneer surface [23]. Therefore, the parameters A and n were first calculated taking into account only those samples which exhibited crack origin on the surface of the veneer layer. No significant difference, however, was observed between these values and those calculated taking into account the total number of specimens. The A and n parameters presented here refer thus to calculations considering all tested specimens, since a larger number of samples allowed better fitting of the experimental data.

Table 3 presents the calculated subcritical crack growth parameters for the veneer material of the investigated composites. Both feldspathic glasses CerS and VM7 exhibited a reasonably high n value, indicating a low susceptibility to subcritical crack growth. The apatite-containing veneer (Eris), on the other hand, presented a surprisingly low value for the exponent n, which reveals a very high susceptiblity towards subcritical crack propagation.

Table 3.

Subcritical crack growth parameters A (A*) and n obtained in this work for the veneer material on the veneer–framework composites submitted to cyclic loading (10 Hz, R = −1)

|  | VM7 in Inc-VM7 | Eris in Emp2-Eris | CerS in TZP-CerS |
| --- | --- | --- | --- |
| A | 7.43 × 10−14 | 2.88 × 10−9 | 4.24 × 10−13 |
| n | 28.2 | 8.4 | 25.4 |
| A* | 2.36 × 10−5 | 9.80 × 10−7 | 1.88 × 10−5 |

The subcritical crack growth behavior of the veneer materials are compared in Fig. 3 with literature data obtained for pure hydroxyapatite and another feldspathic porcelain [21] and [31] using the dynamic fatigue method. In order to compare these results with those obtained in the current study, the literature data were recalculated for the cyclic loading conditions employed in this work, as described elsewhere [24], [25] and [30]. Since in the alternating bending tests the volume of specimen submitted to stress is twice larger than in the dynamic fatigue experiments, a correction factor for the volume effect was also included in these calculations [25].

|  | Display Full Size version of this image (27K) |
| --- | --- |

Fig. 3. Crack velocity data for the subcritical crack growth in the veneer material of the investigated composites, in comparison to literature data for hydroxyapatite and a feldspathic glass [21] and [31].

The crack velocity values calculated in this study for the CerS and VM7 veneers are approximately one order of magnitude lower than the velocity data obtained for the dental porcelain evaluated by Lohbauer et al. [21] (Fig. 3), which might be related to differences in chemical composition and crystalline phases between these materials. The chemical composition is known to have a marked effect on the subcritical crack growth behavior of SiO2 glass-based materials. The presence of alkaline oxides, for instance, considerably increases the water-assisted crack propagation in silicate glasses [32]. In fact, the parameter n calculated for the CerS and VM7 veneers lies in between the values for pure silica (n = 38) and soda-lime glasses (n = 15, 25% alkalis), which can be attributed to the intermediate alkaline content (typically 11%) of dental porcelains [32].

The high susceptibility to subcritical crack growth of the Eris veneer is apparently related to the low resistance of apatite materials against water-assisted crack propagation. Fig. 3 shows that the slope (n) of the crack velocity curve obtained for the fluorapatite-based veneer (Eris) is comparable to that of hydroxyapatite in water of a previous study [31]. The enhanced subcritical crack growth in hydroxyapatite is due to its dissolution into Ca2+, PO43− and OH− ions when submitted to tensile stresses in water [31] and [33]. Fluorapatite is known to be less soluble in water than hydroxyapatite due to the higher crystallinity and density achieved through the substitution of fluor atoms for hydroxyl groups in the apatite lattice [34] and [35]. This might be the reason for the lower crack velocities measured for the fluorapatite-based veneer in comparison to those obtained for pure hydroxyapatite (Fig. 3). However, the lower solubility of fluorapatite is not sufficient to avoid fast subcritical crack growth in water under tensile conditions.

Interestingly, even though it closely resembles the composition of the enamel of natural teeth, hydroxyl- and fluorapatite do not seem to be appropriate materials for biomedical applications where wet and high tensile loading conditions prevail. In fact, the apatite-rich enamel of natural teeth is not exposed to the high tensile stresses that occur on the gingival side of the connector of restorative ceramic bridges. Additionally, the different microstructures of apatite in artificial veneers and in natural teeth certainly also play an important role on the dissolution process and the subcritical crack growth behavior. This reveals the importance of matching the material's properties with the loading conditions, rather than solely mimicking the natural teeth composition, when selecting and developing materials for dental prostheses.

The subcritical crack growth and Weibull parameters obtained here (Table 2 and Table 3) allows us to estimate the lifetime of the veneer layer under different stress conditions, as exemplified in Fig. 4 for a failure probability of 5% [19]. As a result of its high susceptibility to subcritical crack growth, the apatite-based veneer (Eris) exhibits markedly shorter lifetime in comparison to the feldspathic glasses (VM7 and CerS).

|  | Display Full Size version of this image (25K) |
| --- | --- |

Fig. 4. The number of cycles to failure of the veneer layer for different maximum stresses under wet and cyclic loading conditions. The number of cycles was calculated for a failure probability of 5% (Fig. 2). The upper axis displays the expected lifetime for the veneer layer under the wet conditions in the mouth, assuming a mastication frequency of 1400 cycles/day. The inset arrows indicate the maximum stress that can be applied to the different veneer materials in order to achieve a minimum lifetime of 20 years at a failure risk of 5%.

The lifetime diagram shown in Fig. 4 was used to establish guidelines for the selection of veneer materials and for the design of multi-unit dental bridges. According to this graph, in order to extend the lifetime of the veneer layer to more than 20 years at a survival probability of 95% (5% risk of failure), the stress level on the veneer surface should not exceed 15, 40 and 43 MPa for the Eris, CerS and VM7 materials, respectively. Assuming the maximum stress on the bridge to occur on the gingival side of the connectors, one can calculate the minimum connector diameters required to keep the stress level below the limits shown in Fig. 4. These may serve as guidelines for the practitioners when designing a bridge restoration. Using a quite conservative approach, a high chewing frequency of 1400 cycles/day (in Fig. 4) and an occlusal bite force of 250 N during mastication were assumed for the calculations [2]. Additionally, the crack growth parameters shown in Table 3 were assumed to be valid for the entire stress range, despite the fact that a stress intensity factor threshold (KI,th) might exist below which no crack propagation takes place. Even though occlusal biting forces as high as 800 N might be achieved in the molar region [2], this high load level only occurs for very short periods of time and does not represent the long-term loading conditions which determines the subcritical crack growth phenomenon.

The influence of the connector dimensions on the stresses developed in multi-unit bridges was discussed in a recent publication, on the basis of finite element analyses available in the literature and simple calculations assuming free-rotating supported beams [19]. The boundary conditions and further assumptions required for the calculations are described in detail in [19]. As a result, a possible range of stress values on the bridge framework was estimated for round connectors of different diameters.

In the case of the veneer–framework composites evaluated in this study, the lower elastic modulus of the veneer (Eveneer) leads to a reduced stress on this layer in comparison to that developed on the framework core [23]. In this work, the stress on the veneer surface was estimated () using the following equation:

|  | (4) |
| --- | --- |

where hveneer is the thickness of the veneer layer, the previously estimated maximum stress on the framework [19], and Eframe and rframe are the elastic modulus and radius of the framework, respectively. The derivation of this equation is shown in the Appendix A. The contribution of the veneer layer to the overall stiffness of the composite was neglected in these calculations, due to the thinness and low elastic modulus of the veneer coating (see Appendix A). Following the conservative approach, the stress on the gingival side of the framework surface was assumed to be the highest within the stress range estimated in reference [19]. Fig. 5 displays the stresses expected on the veneer surface of various multi-unit bridges for different connector diameters (rframe), assuming a veneer layer thickness (hveneer) of 1 mm in Eq. (4). The elastic modulus data used in these calculations were taken from reference [23].

|  | Display Full Size version of this image (41K) |
| --- | --- |

Fig. 5. Calculated stress on the veneer surface (gingival side) as a function of connector diameter for different veneer materials and types of bridge (see assumptions in the text and in the Appendix A).

Based on Fig. 5, the minimum connector diameter required to reduce the stress on the veneer surface below the critical values indicated in Fig. 4 was estimated. By designing connectors with diameter larger than this minimum value, the risk of veneer failure due to subcritical crack growth is expected to be lower than 5% within a period of 20 years. Table 4 shows that, in general, larger connector sizes are needed to withstand the high mastication stresses on bridges with increasing number of pontics. Thick connectors are specially required in the case of the Eris veneer due to its drastically higher susceptibility to subcritical crack growth (Fig. 3 and Fig. 4). Since restorations in the molar region are to some extent limited in height (4–5 mm) [36], crack propagation from the veneer surface is expected to occur even for shorter 3-unit bridges for the Emp2-Eris material. Cracks propagating from the veneer of Emp2-Eris specimens result in complete failure of the layered composite [23], suggesting that this material cannot be reliably used for posterior restorations.

Table 4.

Minimum diameter (mm) and cross section area (mm2) of the framework connector required to minimize the risk of cracks in the veneer due to subcritical crack growth within a period of 20 yearsa

| Minimum diameter (mm)/minimum cross section area (mm2) |  |  |  |
| --- | --- | --- | --- |
|  | Inc-VM7 | Emp2-Eris | TZP-CerS |
| 3 units | 3.7/10.8 | 6.3/31.2 | 3.9/11.9 |
| 4 units | 5.8/26.4 | 10.1/80.1 | 6.2/30.2 |
| 5 units | 7.0/38.5 | 12.1/115 | 7.4/43.0 |

Connectors with round cross section were assumed in these calculations.
a See assumptions and boundary conditions in the text.

On the other hand, the higher resistance against subcritical crack growth of the feldspathic veneer leads to less stringent connector dimensions for the Inc-VM7 and TZP-CerS composites (Table 4). In the case of short 3-unit bridges, the fracture probability of the veneer under subcritical conditions is markedly low (<5%) for connector diameters equal or larger than 4 mm. For longer 4- and 5-unit bridges, the connector dimensions required to reduce the chance of failure within 20 years are higher than that typically achievable in practice during prosthesis preparation. Therefore, it is quite probable that cracks in the veneer layer occur in 4- and 5-unit posterior bridges within a period of 20 years. However, it is important to note that the fracture of the veneer layer does not lead to complete failure of the bridge in the case of Inc-VM7 and TZP-CerS composites [23], due to crack arresting at the veneer–framework interface. However, cracks in the veneer layer can expose the framework core to an aqueous environment and eventually enhance subcritical crack growth in the framework material. For that reason, the subcritical crack growth of dental framework materials under wet and cyclic conditions is addressed in subsequent publications [19].

- 4. Conclusions

The in vitro lifetime analysis of the veneer–framework composites revealed that water-assisted subcritical crack growth leads to crack propagation from the veneer surface at stress levels appreciably lower than the veneer inert strength. Apatite-based veneers (Eris 2) were observed to be more susceptible to subcritical crack growth than feldspathic veneers (CerS and VM7), probably due to their high dissolution in water under tensile loading conditions.

Based on these in vitro results, Emp2-Eris composites are not appropriate materials for highly-loaded dental bridges in the molar region due to the high susceptibility to subcritical crack growth of the apatite veneer in combination with the absence of crack arresting at the veneer–framework interface. The feldspathic veneers of the TZP-CerS and Inc-VM7 composites, on the other hand, may exhibit lifetimes longer than 20 years if the connector of 3-unit bridges is appropriately designed. In the case of 4- and 5-unit posterior bridges, the connector dimensions required for extended lifetimes are difficult to achieve in practice, which may lead to small cracks in the veneer layer due to subcritical crack growth. However, the veneer cracks in the gingival side of the bridge connector do not result in complete failure of the TZP-CerS and Inc-VM7 composites, due to their ability to stop cracks at the veneer–framework interface. Therefore, the fracture of the veneer only exposes the core framework to the aqueous environment of the mouth. The long-term lifetime of zirconia-based bridges (TZP-CerS and Inc-VM7) is then governed by the water-assisted subcritical crack growth behavior of the framework material, which is the subject of a following publication.

- Acknowledgements

The authors would like to gratefully acknowledge Madeleine Schumacher and Olivier Loeffel (University of Zürich, Center for Dental & Oral Medicine) for the kindness and valuable guidance during preparation of the specimens, as well as Marina Tambasco Studart for the patience and precision while layering most of the specimens evaluated in this study. We are also very thankful to Degudent, Vita Zahnfabrik and Ivoclar Vivadent for kindly supplying their framework and veneering materials for the TZP-CerS, Inc-VM7 and Emp2-Eris composites, respectively.

- References

[1] F. Filser, H. Lüthy, P. Scharer and L. Gauckler, All-ceramic restorations by new direct ceramic machining process (DCM), J Dent Res 77 (1998), p. 762 [Abstract No. 1046].

[2] J.R. Kelly, Ceramics in restorative and prosthetic dentistry, Annu Rev Mater Sci 27 (1997), pp. 443–468. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (83)

[3] H. Lüthy, F. Filser, L. Gauckler and P. Scharer, High reliable zirconia bridges by direct ceramic machining process (DCM), J Dent Res 77 (1998), p. 762 [Abstract No. 1045].

[4] B.R. Lawn, A. Pajares, Y. Zhang, Y. Deng, M.A. Polack and I.K. Lloyd et al., Materials design in the performance of all-ceramic crowns, Biomaterials 25 (2004), pp. 2885–2892. SummaryPlus | Full Text + Links | PDF (275 K) | View Record in Scopus | Cited By in Scopus (27)

[5] W.S. Oh and K.J. Anusavice, Effect of connector design on the fracture resistance of all-ceramic fixed partial dentures, J Prosthet Dent 87 (2002), pp. 536–542. Abstract | PDF (148 K) | View Record in Scopus | Cited By in Scopus (17)

[6] W. Oh, N. Gotzen and K.J. Anusavice, Influence of connector design on fracture probability of ceramic fixed-partial dentures, J Dent Res 81 (2002), pp. 623–627. View Record in Scopus | Cited By in Scopus (19)

[7] J.R. Kelly, J.A. Tesk and J.A. Sorensen, Failure of all-ceramic fixed partial dentures in-vitro and in-vivo—analysis and modeling, J Dent Res 74 (1995), pp. 1253–1258. View Record in Scopus | Cited By in Scopus (93)

[8] F. Filser, P. Kocher, F. Weibel, H. Lüthy, P. Scharer and L.J. Gauckler, Reliability and strength of all-ceramic dental restorations fabricated by direct ceramic machining (DCM), Int J Comput Dent 4 (2001), pp. 89–106. View Record in Scopus | Cited By in Scopus (42)

[9] N. Goetzen, G. Natt, M.M. Morlock and J. Tinschert, Numerical stress and reliability analysis of all-ceramic fixed-partial dentures 2003 Summer Bioengineering Conference, June 25–29, Key Biscayne, Florida (2003), pp. 869–870.

[10] S.O. Koutayas, M. Kern, F. Ferraresso and J.R. Strub, Influence of framework design on fracture strength of mandibular anterior all-ceramic resin-bonded fixed partial dentures, Int J Prosthodont 15 (2002), pp. 223–229. View Record in Scopus | Cited By in Scopus (7)

[11] M. Kern, W.H. Douglas, T. Fechtig, J.R. Strub and R. Delong, Fracture strength of all-porcelain, resin-bonded bridges after testing in an artificial oral environment, J Dent 21 (1993), pp. 117–121. Abstract | View Record in Scopus | Cited By in Scopus (26)

[12] A. Sundh, M. Molin and G. Sjogren, Fracture resistance of yttrium oxide partially-stabilized zirconia all-ceramic bridges after veneering and mechanical fatigue testing, Dent Mater 21 (2005), pp. 476–482. SummaryPlus | Full Text + Links | PDF (174 K) | View Record in Scopus | Cited By in Scopus (16)

[13] H. Lüthy, F. Filser, O. Loeffel, M. Schumacher, L.J. Gauckler and C.H.F. Hammerle, Strength and reliability of four-unit all-ceramic posterior bridges, Dent Mater 21 (2005), pp. 930–937. View Record in Scopus | Cited By in Scopus (13)

[14] T.A. Michalske and S.W. Freiman, A molecular mechanism for stress-corrosion in vitreous silica, J Am Ceram Soc 66 (1983), pp. 284–288. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (125)

[15] J. Chevalier, C. Olagnon and G. Fantozzi, Subcritical crack propagation in 3Y-TZP ceramics: static and cyclic fatigue, J Am Ceram Soc 82 (1999), pp. 3129–3138. View Record in Scopus | Cited By in Scopus (33)

[16] D.S. Jacobs and I.W. Chen, Cyclic fatigue in ceramics—a balance between crack shielding accumulation and degradation, J Am Ceram Soc 78 (1995), pp. 513–520. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (36)

[17] R.H. Dauskardt, W. Yu and R.O. Ritchie, Fatigue crack-propagation in transformation-toughened zirconia ceramic, J Am Ceram Soc 70 (1987), pp. C248–C252.

[18] T. Fett, D. Munz and G. Thun, Influence of frequency on cyclic fatigue of coarse-grained Al2o3, J Mater Sci Lett 12 (1993), pp. 220–222. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (7)

[19] A.R. Studart, F. Filser, P. Kocher and L.J. Gauckler, Fatigue of zirconia under cyclic loading in water and its implications for the design of dental bridges, Dent Mater 23 (2007), pp. 106–114. SummaryPlus | Full Text + Links | PDF (545 K) | View Record in Scopus | Cited By in Scopus (4)

[20] H. Fischer, M. Weber and R. Marx, Lifetime prediction of all-ceramic bridges by computational methods, J Dent Res 82 (2003), pp. 238–242. View Record in Scopus | Cited By in Scopus (20)

[21] U. Lohbauer, A. Petschelt and P. Greil, Lifetime prediction of CAD/CAM dental ceramics, J Biomed Mater Res 63 (2002), pp. 780–785. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (13)

[22] J. Geis-Gerstorfer, P. Fassler and R. Kirmeier, Fatigue behavior of three all ceramic materials, J Dent Res 81 (2002), p. A469 [Abstract No. 3835].

[23] A.R. Studart, F. Filser, P. Kocher, H. Lüthy and L.J. Gauckler, Mechanical and fracture behavior of veneer–framework composites for all-ceramic dental bridges, Dent Mater 23 (2007), pp. 115–123. SummaryPlus | Full Text + Links | PDF (577 K) | View Record in Scopus | Cited By in Scopus (2)

[24] A.G. Evans and E.R. Fuller, Crack-propagation in ceramic materials under cyclic loading conditions, Metall Trans 5 (1974), pp. 27–33. View Record in Scopus | Cited By in Scopus (98)

[25] D. Munz and T. Fett, Ceramics: mechanical properties, failure behaviour, materials selection, Springer-Verlag, Berlin (1999).

[26] V. Imbeni, J.J. Kruzic, G.W. Marshall, S.J. Marshall and R.O. Ritchie, The dentin–enamel junction and the fracture of human teeth, Nat Mater 4 (2005), pp. 229–232. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (20)

[27] P. Vanlanduyt, F. Li, J.P. Keustermans, J.M. Streydio, F. Delannay and E. Munting, The influence of high sintering temperatures on the mechanical-properties of hydroxylapatite, J Mater Sci-Mater Med 6 (1995), pp. 8–13. View Record in Scopus | Cited By in Scopus (72)

[28] M. Guazzatto, K. Proos, L. Quach and M.V. Swain, Strength, reliability and mode of fracture of bilayered porcelain/zirconia (Y-TZP) dental ceramics, Biomaterials 25 (2004), pp. 5045–5052.

[29] R. Schmitt, T. Fett and D. Munz, Cyclic fatigue of zirconia, Fatigue Fract Eng Mater Struct 19 (1996), pp. 1411–1420. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (7)

[30] T. Fett, G. Martin, D. Munz and G. Thun, Determination of Da/Dn-Delta-K1 curves for small cracks in alumina in alternating bending tests, J Mater Sci 26 (1991), pp. 3320–3328. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (19)

[31] S. Raynaud, E. Champion, D. Bernache-Assolant and D. Tetard, Dynamic fatigue and degradation in solution of hydroxyapatite ceramics, J Mater Sci-Mater Med 9 (1998), pp. 221–227. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (18)

[32] S.M. Wiederhorn and L.H. Bolz, Stress corrosion and static fatigue of glass, J Am Ceram Soc 53 (1970), p. 543. Full Text via CrossRef

[33] S.V. Dorozhkin, A review on the dissolution models of calcium apatites, Prog Cryst Growth Charact Mater 44 (2002), pp. 45–61. Abstract | Abstract + References | PDF (1361 K) | View Record in Scopus | Cited By in Scopus (19)

[34] Y.M. Chen and X.G. Miao, Effect of fluorine addition on the corrosion resistance of hydroxyapatite ceramics, Ceram Int 30 (2004), pp. 1961–1965. SummaryPlus | Full Text + Links | PDF (360 K) | View Record in Scopus | Cited By in Scopus (7)

[35] N.H. de Leeuw, Resisting the onset of hydroxyapatite dissolution through the incorporation of fluoride, J Phys Chem B 108 (2004), pp. 1809–1811. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (11)

[36] J.F. Esquivel-Upshaw, K.J. Anusavice, H. Young, J. Jones and C. Gibbs, Clinical performance of a lithia disilicate-based core ceramic for three-unit posterior FPDs, Int J Prosthodont 17 (2004), pp. 469–475. View Record in Scopus | Cited By in Scopus (12)

[37] M. Beghini, L. Bertini and F. Frendo, Measurement of coatings’ elastic properties by mechanical methods: part 1. Consideration on experimental errors, Exp Mech 41 (2001), pp. 293–304. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (4)

[38] S.P. Timoshenko and J.M. Gere, Mechanics of materials, Litton Educational Publishing, Inc., London (1972).

- Appendix A.

The stress on the veneer surface of multi-unit bridges at the gingival side of the connector was estimated using beam theory calculations assuming free-rotating supported beams. A cylindrical beam displaying the same diameter as a round bridge connector was taken as a representative model. Fig. A.1 schematically shows the stress and strain profiles along the z direction (y = 0) for the model composite beams considered in the calculations. The centroid (neutral axis) of the composite beam coincides in this case with the specimen geometric center, leading to a symmetric (but oppositely signed) stress and strain distribution (Fig. A.1).

|  | Display Full Size version of this image (29K) |
| --- | --- |

Fig. A.1. Stress (σ) and strain () distribution along the z direction (y = 0) for a model cylindrical beam consisting of a framework core (in white) and an outer veneer layer (in grey). In this case, the bending moment M is applied in the xz plane.

Based on beam theory, the longitudinal strain (x) along the model composite beam for an applied moment M in the xz plane is given by [37] and [38]:

|  | (A.1) |
| --- | --- |

where Ek and Jk are the elastic modulus and the moment of inertia of the kth component of the composite beam, respectively (in this case, the veneer layer and the framework core). Taking into account that the elastic modulus of the framework (Eframe) is considerably higher than that of the veneer material (Eveneer), the contribution of the veneer coating to the overall stiffness () was neglected in the calculations, leading us to the following simplified form of Eq. (A.1):

|  | (A.2) |
| --- | --- |

The strain at the veneer–framework interface (int) is then given by Eq. (A.2) for z = rframe, whereas the strain on the veneer surface (surf) is obtained when z = (rframe + hveneer). From these strain values, the stress along the composite beam can be calculated using Hooke's law (σ = E):

|  | (A.3) |
| --- | --- |

|  | (A.4) |
| --- | --- |

|  | (A.5) |
| --- | --- |

where , and are the stress values indicated in Fig. A.1.

By substituting the appropriate strain values from Eq. (A.2) into Eqs. (A.3) and (A.5), and subsequently dividing Eq. (A.5) by Eq. (A.3), the following relation can be obtained for the stress on the veneer surface:

|  | (A.6) |
| --- | --- |

where rframe and hveneer are the radius of the framework core and the thickness of the veneer layer, respectively (Fig. A.1). The stress on the veneer surface was estimated from this equation, based on the maximum stress values calculated on a previous study [19] for the framework core ().

Correspondence to: Department of Materials, Nonmetallic Inorganic Materials, Wolfgang–Pauli-Strasse 10, ETH Hönggerberg, HCI G 539, 8093 Zürich, Switzerland. Tel.: +41 44 632 37 18; fax: +41 44 632 11 32.

Dental Materials Volume 23, Issue 2, February 2007, Pages 177-185

Cíclicos fadiga em água do verniz - quadro compósitos para todos - cerâmica dentária pontes

André R. Studarta,, Frank Filsera, Peter Kochera, Heinz Lüthyb e Ludwig J. Gaucklera 
ADepartment dos Materiais, ETH - Zürich, Zurique, Suíça 
BDental School, da Universidade de Basel, Basel, Suíça 
Recebido 3 de Agosto de 2005; Aceite 9 de Janeiro de 2006. Disponível em linha 21 de Fevereiro de 2006.

Abstract 
Objectivos 
As cerâmicas aplicadas na odontologia podem apresentar um crescimento das rachaduras subcríticas devido à grave carregamento cíclico no meio aquoso encontradas na boca durante a mastigação.

Avaliaram o comportamento de três sistemas de restauração dentária (Empress 2 / IPS Eris, TZP / Cercon S e Inceram-Zirconia/Vita VM7) sob carregamento cíclico em água, a fim de estabelecer diretrizes para o uso e concepção de vida longa as cerâmicas de pontes posteriores.

Métodos 
Forças inertes e testes sob carregamento cíclico foram realizados em um ambiente aquoso. Foi utilizado um aparelho de flexão mecânica e avaliadas com as estatísticas de Weibull.

Resultados 
Ocorreu um crescimento das rachaduras subcríticas predominantemente no exterior da camada de verniz. O verniz a base de apatita (IPS Eris) era mais suscetível à propagação de trincas subcríticas, em comparação com as coroas de cerâmica feldspatica (Cercon S e Vita VM7).

Significância 
Os sistemas de restauração dental contendo apatita (Empress 2 / IPS Eris) não são recomendados para a fabricação de pontes de cerâmicas na região de molares. No entanto as cerêmicas feldspaticas e Zircônia infiltrada por vidro (TZP / Cercon S e Inceram-Zirconia/Vita VM7) podem apresentar vida útil superior a 20 anos, se a ponte e o conector estão devidamente projetados.

Palavras - chave: Dental ponte; Corrigido parcial denture; Cerâmica; Zirconia; Alumina; Porcelana; Cíclicos fadiga; Subcríticos rachadura crescimento; Lifetime 

Artigo Outline 
1. Introdução 
2. Materiais e métodos 
3. Resultados e discussão 
4. Conclusões 
Agradecimentos 
Apêndice A. Apêndice 
Referências 

1. Introdução 
Layered cerâmicos compósitos foram amplamente utilizado nos últimos anos como único coroas e multi - unidade pontes em odontologia reparadora. A substituição da tradicional metalúrgica de base restaurações com todos - cerâmicos próteses tem sido impulsionada pela melhoria estética e menor potencial alergénio realizados utilizando materiais cerâmicos como núcleo quadros [1], [2] e [3]. 
Contudo, a natureza frágil e tipicamente baixa confiabilidade dos materiais cerâmicos exige uma análise profunda do comportamento mecânico e fadiga de cerâmica restaurações, a fim de garantir longa duração pontes e coroas. Extensos estudos têm sido levadas a cabo pelos relvados e Al. [4] sobre as propriedades mecânicas e fadiga de dentista cerâmica sob o desgaste típico e entre em contato com carga condições aplicadas a mergulhados coroas. 
Além de desgaste e contato carga e multi - unidade pontes são também submetidos à alta tensão insiste no gengival lado do conector durante a mastigação. Desde cerâmicos são muito sensíveis à falência sob tracção carregamento, todos - cerâmicos pontes exigir ainda mais rigorosas do que as propriedades mecânicas necessárias para coroas dentais. Portanto, vários autores têm recentemente investigou a resistência mecânica e capacidade de carga ostentando todos - cerâmicos pontes [5], [6], [7], [8], [9], [10], [11], [12] E [13]. De acordo com estes estudos, com base zircónia - cerâmicas são os mais adequados para suportar a alta tensão salienta que ocorrem em múltiplos unidade pontes. 
Além da alta resistência mecânica inicial, um aspecto fundamental que tem de ser considerado quando se utiliza cerâmica como componentes estruturais é a sua susceptibilidade a trinca subcrítico crescimento fenómenos. Em geral, o crescimento subcrítico rachadura em materiais cerâmicos é causada pela água - assistida quebra de metal - óxido de títulos na ponta da trinca sob estresse aplicado [14]. A propagação rachadura reduz a resistência do componente cerâmica, levando ao fracasso em salientar as condições em que nenhuma fratura seria de esperar, com base unicamente no primeiro inertes força. No caso das cerâmicas mais rígida, cíclico carga aumenta a água - assistida subcríticos rachadura propagação devido à degradação do material da toughening mecanismos [15], [16], [17] e [18]. O crescimento subcrítico rachadura comportamento de materiais cerâmicos sob carregamento cíclico pode ser descrita pela seguinte equação poder - lei: 

  
(1) 
Onde v é a velocidade rachadura, Δ KI a amplitude do estresse intensidade fator aplicado (KI, KI - no máximo, minutos), KI, max e KI, min, a intensidade máxima e mínima estresse fatores, o estresse R rácio KI, minutos / KI , No máximo, e A e n são os parâmetros de crescimento subcrítico rachadura carregamento cíclico. 
O ambiente aquoso da boca e cíclico carregamento imparted durante a mastigação são aparentemente muito favoráveis condições para subcríticos rachadura propagação restaurações em cerâmica. A avaliação do comportamento do crescimento subcrítico rachadura dentária cerâmica em tais condições, portanto, é altamente desejável [2] e [19]. Estudos anteriores incidiram sobre as propriedades do verniz e quadro materiais separadamente, dando informações valiosas sobre o crescimento rachadura parâmetros dos componentes individuais da restauração dentária [19], [20], [21] e [22]. 
No entanto, a estimativa da vida de uma prótese dentária também requer o conhecimento sobre o crescimento subcrítico rachadura comportamento do verniz e quadro materiais quando montados juntos na final mergulhados compósito. Investigações sobre o verniz - quadro compósito prestar informações importantes sobre o estresse distribuição, o fracasso modo, rachadura origem, interfacial caução e térmica salienta, que não pode ser avaliado através da avaliação dos materiais separadamente. 
Um artigo anterior sobre o rápido fratura (KI, aplicada = KIC) camadas de compósitos [23] mostrou que originam em fissuras, ou perto, o verniz superficial e propagadas para o verniz - quadro Interface. A rachadura origem é definida pela distribuição do estresse sobre o compósito e as falha população introduzida no material durante o processamento. 
Sob condições subcríticos (KI, aplicados <KIC), o tamanho inicial falha na distribuição do verniz e camadas quadro muda em função do tempo, de acordo com o crescimento subcrítico rachadura comportamento de cada material. Isso pode levar a falência diferentes modos e rachadura origens do que as observadas em rápido fratura. Em restaurações dentárias, a trinca subcrítico propagação no verniz camada Espera - se que seja reforçada pelo contacto directo deste material com a boca do ambiente aquoso. Carregamento cíclico, por outro lado, não é suposto a favorecer o crescimento rachadura no verniz camada devido à glassy, não mais rígida natureza desse material [24] e [25]. Em contrapartida, o quadro mais rígida núcleo é mais susceptível à toughening degradação causada por mecanismos cíclico carregamento [15], [16], [17] e [18]. No entanto, não significativo água - assistida rachadura crescimento é esperada para ocorrer no quadro, enquanto o verniz camada protege o núcleo material de exposição direta à água. Estes efeitos concomitantes devem ser considerados ao estimar o tempo de vida de todos - cerâmicos pontes, mas até ao momento não tenham sido abordadas na literatura. 
O objetivo deste trabalho é investigar a vida eo crescimento subcrítico rachadura comportamento de três verniz - quadro dentária compósitos que têm sido utilizados ou considerados como potenciais candidatos para todos - cerâmica dentária pontes. Lifetime experimentos foram conduzidos em água e sob carregamento cíclico, simulando as condições encontradas na boca durante a mastigação. Retangular mergulhados amostras foram utilizadas para controlar melhor a camada de espessura durante a preparação e para avaliar o salienta desenvolvidos sobre o compósito usando simples feixe teoria cálculos. Baseado no subcríticos rachadura crescimento comportamento dessas camadas compósitos, orientações podem ser obtidas para a adequada concepção e selecção de verniz e quadro materiais cerâmicos para todos - restaurações. 
2. Materiais e métodos 
O verniz - quadro compósitos investigados neste estudo são apresentados na Tabela 1. As camadas espécimes (4 milímetros × 5 milímetros x 50 milímetros) foram elaboradas na sequência do multi-layering/firing medidas recomendadas pelos fornecedores e descrito em detalhes em uma recente publicação [23]. 
Tabela 1. 
Verniz - quadro compósitos investigados neste estudo 
________________________________________ Inc - VM7 
________________________________________ Emp2 - Eris 
________________________________________ TZP - CerS 
________________________________________ 
Framework materiais Glass infiltradas Al2O3 - ZrO2 compósito (Inceram - Zirconia) Li2O.2SiO2 vidro - cerâmica (Empress 2) Yttria - estabilizado ZrO2 (3Y - TZP, Cercon) 
Verniz materiais Feldspathic vidro (VITA VM 7 sistema) Fluorapatite vidro - cerâmica (IPS Eris para E2 sistema) Feldspathic vidro (Cercon Ceram S) 
Supplier VITA Zahnfabrik, Alemanha Ivoclar Vivadent, Liechtenstein Degudent, Alemanha 
Para determinar o crescimento subcrítico rachadura comportamento do verniz - quadro compósitos, aproximadamente 30 camadas amostras de cada composição foram avaliados em termos de duração e inerte força, segundo a metodologia descrita em um artigo anterior [19]. 
A força inerte medições foram realizadas em flexão modo, a aplicação de uma taxa de carregamento rápido (> 102 MPa S - 1), a fim de evitar subcríticos rachadura crescimento durante o experimento. Neste teste rápido fratura, o espécime for pela primeira vez estabelecida entre uma parede rígida e uma longa cantilever. A flexão momento é então aplicada ao modelo por aumentar continuamente a carga no final do cantilever. 
A vida útil testes foram realizados através do envio das amostras para um sinusoidais cíclico carga e gravação do tempo necessários para o fracasso. Cíclicos carregamento foi aplicado de um modo alternado (rácio R = σ minutos / σ máx = -1) com uma frequência de 10 Hz, mantendo uma constante máxima / mínima estresse sobre o verniz superficial. Apesar da resistência dos inertes semelhante folheados, os ensaios foram realizados em um nível máximo de 50 MPa estresse para o Emp2 - Eris amostras, e 65 MPa para o TZP - CerS e Inc - VM7 espécimes. Estes níveis correspondem a uma tensão aplicada momento de, aproximadamente, 700, 1030 e 1070 N milímetros no Emp2 - Eris, TZP - CerS e Inc - VM7 compósitos, respectivamente. A magnitude da tensão máxima foi escolhida com base em testes preliminares, a partir do qual o estresse nível necessário para induzir amostra insucesso dentro de um montante razoável de tempo (<11 dias) foi criado. 
Mais detalhes sobre a flexão máquina e testar as configurações utilizadas para medições acima podem ser encontrados na referência [19]. O inertes força e vida dados obtidos por este meio foram posteriormente avaliadas utilizando Weibull estatísticas. 
3. Resultados e discussão 
A fratura do mergulhados compósitos sob carregamento rápido condições naturais foi iniciado a partir de rachaduras no verniz camada que propagadas para o verniz - quadro da relação, como discutido num anterior inquérito [23]. Em compósitos contendo um núcleo duro quadro (TZP - CerS e Inc - VM7), a trinca propagação da camada verniz foi desviar na interface semelhante ao rachadura prender mecanismo que ocorre na junção dentina - esmalte de dentes naturais [26]. Fig. 1 mostra a tensão necessária para reprimir propagação do verniz camada estimado de feixe teoria cálculos [23]. A característica força (σ c, 0) e Weibull modulo (metros), do folheados, não variam consideravelmente entre os investigados compósitos (Tabela 2), com valores dentro do tipicamente baixo nível de resistência porcelana óculos. 
  
Display Full Size versão desta imagem (21K) 

Fig. 1. Weibull distribuição de inertes força da camada de verniz o verniz - quadro compósitos, de acordo com Studart e Al. [23]. 
Tabela 2. 
Weibull parâmetros obtidos para os inertes força e vida útil do material de verniz mergulhados espécimes 
________________________________________ VM7 em Inc - VM7 
________________________________________ Eris em Emp2 - Eris 
________________________________________ CerS em TZP - CerS 
________________________________________ 
Inertes força 
  Metros 8,9 9,5 9,2 
  Σ c, 0 (MPa) 107,4 96,6 103,9 

Lifetime 
  Σ máx aplicada (MPa) 65 50 65 
  Metros * 0,34 1,48 0,39 
  Nf, 0 (ciclos) 490050 3210 73690 
A vida útil experimentos foram realizados através da aplicação de uma carga cíclica dobra com estresse máximo fixado para o verniz superficial. Todas as amostras testadas sob húmido cíclico carregamento fraturado originários de rachaduras no verniz camada exterior. A precisão rachadura origem não foi tão evidente como no rápido fraturado amostras, devido ao menor estresse grandeza aplicados na vida experimentos. Estes resultados indicam que a trinca subcrítico propagação, em condições similares às encontradas na boca leva à fratura do verniz camada antes da ruptura do núcleo quadro materiais. A água - assistida fratura do verniz materiais, portanto, ocorre mais rapidamente do que a carga cíclica - induzida ruptura do interior quadro materiais. 
Lifetime dados obtidos para o verniz - quadro compósitos sob carregamento cíclico em água são mostrados na Fig. 2. Apesar do menor tensão máxima aplicada (50 MPa), o Emp2 - Eris compósito mostrou uma surpreendentemente baixa resistência à trinca subcrítico crescimento, apresentando vida útil significativamente mais curtos do que os obtidos para o TZP - CerS e Inc - VM7 espécimes. A fratura do Emp2 - Eris amostras ocorreu no mesmo modo que o observado no carregamento rápido experimentos [23], com a propagação da trinca verniz camada em todo o compósito sem qualquer desvio rachadura no verniz - quadro Interface. A vida útil da Emp2 - Eris compósitos foi, assim, exclusivamente determinado pelo crescimento subcrítico rachadura comportamento do verniz materiais, o que indica que estas restaurações são altamente sensíveis à completa falta bastante baixa tensão em condições (50 MPa). 
  
Display Full Size versão desta imagem (25K) 

Fig. 2. Weibull distribuição para o número de ciclos de fracasso da camada de verniz o verniz - quadro compósitos. 
No zircónia - baseado compósitos (TZP - CerS e Inc - VM7), por outro lado, rachaduras foram deflectidas no verniz - quadro interface como na fratura testes rápidos [23]. Devido ao desvio rachadura na interface, as primeiras fissuras na camada de verniz Inc - VM7 e TZP - CerS compósitos não levam a concluir fracasso do todo - cerâmicos restauração, mas apenas expõe o quadro central para o ambiente aquoso. A vida útil de dados mostrado na figura. 2 referem - se apenas à fratura do verniz materiais, uma vez que o estresse nível aplicado não foi alto o suficiente para quebrar o zircónia - quadro baseado núcleo após rachadura desvio na interface. A vida do núcleo quadro sob a aquosa e cíclico condições esperada após o fracasso do verniz camada será abordada numa sequência da publicação. 
A camada de verniz a TZP - CerS e Inc - VM7 compósitos exibiu vidas acentuadamente mais do que o obtido para o Emp2 - Eris espécimes, apesar da maior tensão máxima aplicada durante o ensaio (65 MPa). No caso do Inc - VM7 compósito, a camada de verniz alguns dos espécimes não rachadura, mesmo dentro de 11 dias de prova (Fig. 2). Uma vez que a probabilidade de insucesso devido ao crescimento subcrítico rachadura é determinada pela população inicial falha no material (neste caso, o verniz), a vida útil de dados também pode ser descrita em termos de Weibull estatísticas [19], como mostrado na figura. 2 andTable 2. 
O Weibull vida dados descritos na Fig. 2 e os inertes força resultados demonstrados na Fig. 1 podem ser combinadas para derivar o crescimento subcrítico rachadura parâmetros A e n do verniz materiais (Eq. (1)), usando as seguintes equações [19] e [25]: 

  
(2) 
E 

  
(3) 
Onde milhões * e Nf, 0 são os Weibull módulo e as características número de ciclos de fracasso da vida distribuição (Fig. 2 e Quadro 2), KIC o estresse intensidade fator crítico, a Y geométrica constante relacionada com a localização e forma rachadura (Equivalente a 1,3 por fissuras superficiais) e Δσ é a amplitude da tensão aplicada. Critical estresse intensidade factores (KIC) relatados na literatura para apatite (1,0 MPa m1 / 2) [27] e para feldspathic porcelains (0,9 MPa m1 / 2) [28] foram utilizados para calcular os parâmetros A e n Para a Eris, e os VM7 e CerS folheados, respectivamente. 
É importante notar que os parâmetros A e n só pode ser calculado a partir Eqs. (2) e (3), quando as fraturas obtidas no inertes força ea vida medições são causados pela mesma falha população no interior do material [18], [25], [29] e [30]. Mesmo que não foi possível identificar a origem exacta das rachaduras em espécimes submetidos aos testes a vida, é razoável supor que, neste caso, as primeiras rachaduras propagados do verniz superficial onde o stress é mais elevado eo material está directamente exposta à Água. Pela força medições, por outro lado, apenas uma parte (40-60%) dos espécimes fraturados de falhas no verniz superficial [23]. Portanto, os parâmetros A e n foram inicialmente calculadas tendo em conta apenas as amostras que exibiu rachadura origem na superfície da camada verniz. Nenhuma diferença significativa, no entanto, foi observada entre estes e os valores calculados tendo em conta o número total de amostras. A UM e n parâmetros aqui apresentados referem, assim, para cálculos considerando todas testadas amostras, uma vez que um maior número de amostras permitiu melhor montagem dos dados experimentais. 
Tabela 3 apresenta o cálculo subcríticos rachadura crescimento parâmetros para o verniz dos materiais compósitos investigados. Ambos feldspathic óculos CerS e VM7 exibiu uma n. valor razoavelmente elevado, indicando uma baixa susceptibilidade ao crescimento subcrítico rachadura. A apatite - contendo verniz (Eris), por outro lado, apresentou uma surpreendentemente baixo valor para o expoente n., o que revela uma muito elevada susceptiblity para subcríticos rachadura propagação. 
Tabela 3. 
Subcríticos rachadura crescimento parâmetros A (A *) e n obtidos neste trabalho para o verniz sobre o verniz materiais compósitos - quadro apresentado ao cíclico carregamento (10 Hz, R = -1) 
________________________________________ VM7 em Inc - VM7 
________________________________________ Eris em Emp2 - Eris 
________________________________________ CerS em TZP - CerS 
________________________________________ 
A 7,43 × 10-14 2,88 × 10-9 4,24 × 10-13 
N. 28,2 8,4 25,4 
A * 2,36 × 10-5 9,80 × 10-7 1,88 × 10-5 
O crescimento subcrítico rachadura comportamento do verniz materiais são comparados em Fig. 3 com literatura dados obtidos por pura hidroxiapatita e outro feldspathic porcelana [21] e [31] utilizando o método dinâmico fadiga. A fim de comparar os resultados com os obtidos no presente estudo, a literatura dados foram recalculados para o carregamento cíclico condições empregadas neste trabalho, como descrito noutros países [24], [25] e [30]. Desde no alternando vincos testes, o volume de amostra submetida ao estresse é duas vezes maior do que na dinâmica fadiga experiências, de um factor de correcção para o volume efeito também foi incluído nestes cálculos [25]. 
  
Display Full Size versão desta imagem (27K) 

Fig. 3. Crack velocidade de dados para o crescimento subcrítico rachadura no verniz dos materiais compósitos investigados, em comparação com os dados de literatura hidroxiapatita e um feldspathic vidro [21] e [31]. 
A rachadura velocidade valores calculados neste estudo para a CerS e VM7 folheados são aproximadamente uma ordem de grandeza mais baixa do que a velocidade dos dados obtidos para as porcelanas dentárias avaliadas por Lohbauer e Al. [21] (Fig. 3), o que pode estar relacionado com diferenças de composição química e cristalinas fases entre esses materiais. A composição química é conhecida por ter um efeito marcado sobre o crescimento subcrítico rachadura de SiO2 vidro comportamento materiais de base. A presença de óxidos alcalino, por exemplo, aumenta consideravelmente a água - assistida rachadura na propagação silicato óculos [32]. Com efeito, o parâmetro n. calculado para o CerS e VM7 folheados reside na entre os valores de pura sílica (N = 38) e carbonato de cal óculos (N = 15, 25% álcalis), o que pode ser atribuído à intermédio alcalina conteúdo (Normalmente 11%) de dentista porcelains [32]. 
A elevada susceptibilidade à trinca subcrítico crescimento da Eris verniz está aparentemente relacionado à baixa resistência dos materiais apatite contra a água - assistida rachadura propagação. Fig. 3 mostra que o declive (n) da velocidade da trinca curva obtida para a fluorapatite - baseado verniz (Eris) é comparável à de hidroxiapatita na água de um estudo anterior [31]. A maior crescimento subcrítico rachadura na hidroxiapatita é devido à sua dissolução em Ca2 +, PO43 - e OH - íons quando submetidos à tracção insiste em água [31] e [33]. Fluorapatite é conhecido por ser menos solúvel em água do que hidroxiapatita devido à maior crystallinity e densidade alcançado através da substituição de fluor átomos de grupos hidroxilo no apatite malha [34] e [35]. Esta poderá ser a razão para a menor fissura velocidades medido para o fluorapatite - baseado verniz, em comparação com os obtidos por pura hidroxiapatita (Fig. 3). No entanto, a menor solubilidade do fluorapatite não é suficiente para evitar o crescimento rápido subcríticos rachadura na água sob tensão condições. 
Curiosamente, embora estreitamente assemelha a composição do esmalte de dentes naturais, hidroxilo - e fluorapatite não parecem ser adequados materiais para aplicações biomédicas onde molhada e alta tensão carregamento condições que prevalecem. De facto, a apatite - ricos esmalte dos dentes naturais não é exposta à alta tensão salienta que ocorrem na região gengival lado do conector de cerâmica reparadora pontes. Além disso, as diferentes microestruturas de apatite na artificiais e folheados em dentes naturais certamente também desempenhar um papel importante no processo de dissolução da trinca subcrítico crescimento comportamento. Isto revela a importância da adequação do material de propriedades com as condições de carga, em vez de apenas imitando os dentes naturais composição, ao selecionar e desenvolver materiais para próteses dentárias. 
A trinca subcrítico crescimento e Weibull parâmetros obtidos aqui (Quadro 2 e Quadro 3) nos permite estimar o tempo de vida do verniz camada sob diferentes condições de estresse, como exemplificado na figura. 4 para uma probabilidade de 5% insucesso [19]. Como resultado de sua elevada susceptibilidade à trinca subcrítico crescimento, a apatite - baseado verniz (Eris) exibe acentuadamente menor duração, em comparação com a feldspathic óculos (VM7 e CerS). 
  
Display Full Size versão desta imagem (25K) 

Fig. 4. O número de ciclos de fracasso da camada de verniz diferentes máximo insiste em húmido cíclico e carregamento condições. O número de ciclos foi calculada para um fracasso probabilidade de 5% (Fig. 2). A parte superior do eixo exibe a vida útil esperada para o verniz camada sob a condições molhadas na boca, assumindo uma mastigação freqüência de 1400 ciclos / dia. O inserida setas indicam a máxima tensão que pode ser aplicado às diferentes verniz materiais, a fim de alcançar uma duração mínima de 20 anos em uma falha risco de 5%. 
A vida útil diagrama mostrado na figura. 4 foi utilizado para estabelecer diretrizes para a seleção de verniz e materiais necessários para a construção de pontes múltiplas unidade odontológica. De acordo com este gráfico, a fim de prolongar o tempo de vida do verniz camada de mais de 20 anos com uma probabilidade de sobrevivência de 95% (5% risco de fracasso), o estresse nível sobre o verniz superfície não deve exceder os 15, 40 e 43 MPa para o Eris, CerS e VM7 materiais, respectivamente. Assumindo a máxima tensão sobre a ponte para ocorrer no gengival lado dos conectores, pode - se calcular o conector diâmetros mínimo necessário para manter o estresse nível abaixo dos limites indicados na figura. 4. Essas podem servir como diretrizes para a praticantes ao projetar uma ponte restauração. Utilizando uma abordagem bastante conservadora, um elevado roer frequência de 1400 ciclos / dia (em Fig. 4) e uma mordida oclusal força de 250 N durante a mastigação foram assumidas para os cálculos [2]. Além disso, a trinca crescimento parâmetros apresentados na Tabela 3 foram assumidos a ser válida para toda a gama frisar, não obstante o facto de um estresse intensidade fator limiar (KI, ª) abaixo do qual não poderiam existir rachadura propagação ocorre. Apesar de mordida oclusal forças tão elevadas como 800 N pode ser alcançada na região molares [2], este elevado nível de carga ocorre apenas por períodos de tempo muito curto e que não representam a longo prazo carregamento condições que determina a trinca subcrítico crescimento fenômeno . 
A influência do conector dimensões no salienta desenvolvido em múltiplos unidade pontes foi discutido em uma recente publicação, com base em elementos finitos análises disponíveis na literatura e simples cálculos assumindo livre - rotativa apoiadas vigas [19]. Os limites e condições mais pressupostos necessários para os cálculos são descritos em detalhes em [19]. Como resultado, uma gama possível de estresse valores ponte sobre o quadro foi estimado por volta de conectores diferentes diâmetros. 
No caso do verniz - quadro compósitos avaliados neste estudo, o menor módulo elástico do verniz (Eveneer) leva a uma redução do stress sobre esta camada, em comparação com o que se desenvolveram no quadro núcleo [23]. Neste trabalho, a tensão sobre o verniz superfície foi estimado (), utilizando a seguinte equação: 

  
(4) 
Onde hveneer é a espessura da camada verniz, o anteriormente estimado máximo estresse sobre o quadro [19], e Eframe e rframe são o modulo elástica e raio do quadro, respectivamente. A derivação desta equação é apresentado no Apêndice A. A contribuição da camada de verniz global da rigidez do compósito foi negligenciada nestes cálculos, devido à magreza e baixa elástica módulo do verniz revestimento (ver Apêndice A). Na sequência da abordagem conservadora, o stress no gengival lado do quadro superfície foi considerada como a mais alta dentro do estresse intervalo estimado na referência [19]. Fig. 5 exibe o salienta esperado sobre o verniz da superfície das diversas unidades múltiplas pontes de diferentes diâmetros conector (rframe), assumindo um verniz camada espessura (hveneer) de 1 mm em Eq. (4). O módulo elástico dados utilizados nestes cálculos foram retirados de referência [23]. 
  
Display Full Size versão desta imagem (41K) 

Fig. 5. Calculada sobre o estresse verniz superficial (gengival lado), em função do diâmetro conector para diferentes tipos de materiais e verniz ponte (ver hipóteses no texto e no Apêndice A). 
Baseada em Fig. 5, o conector diâmetro mínimo necessário para reduzir a tensão sobre o verniz superficial abaixo da crítica dos valores indicados no Fig. 4 foi estimado. Ao projetar conectores com diâmetro maior do que esse valor mínimo, o risco de verniz fracasso devido ao crescimento subcrítico rachadura se espera venha a ser inferior a 5%, num prazo de 20 anos. Quadro 4 mostra que, em geral, maiores conector tamanhos são necessários para suportar o elevado mastigação insiste em pontes com o número crescente de pontics. Thick conectores são especialmente necessária no caso dos Eris verniz drasticamente devido à sua maior susceptibilidade a trinca subcrítico crescimento (Fig. 3 e Fig. 4). Desde restaurações em molares região são, em certa medida limitada em altura (4-5 milímetros) [36], rachadura propagação do verniz superficial é esperada para ocorrer mesmo para curtas 3 - unidade pontes para o Emp2 - Eris materiais. Rachaduras propagação do verniz de Emp2 - Eris espécimes resultar em completo fracasso das camadas compostas [23], o que sugere que esse material não pode ser confiavelmente utilizado para restaurações posteriores. 
Tabela 4. 
Diâmetro mínimo (mm) e seção transversal área (mm2) do quadro conector necessárias para minimizar o risco de rachaduras no verniz devido ao crescimento subcrítico rachadura no prazo de 20 yearsa 
Diâmetro mínimo (mês) / mínimo secção transversal área (mm2) 
________________________________________ 
________________________________________ Inc - VM7 
________________________________________ Emp2 - Eris 
________________________________________ TZP - CerS 
________________________________________ 
3 unidades 3.7/10.8 6.3/31.2 3.9/11.9 
4 unidades 5.8/26.4 10.1/80.1 6.2/30.2 
5 unidades 7.0/38.5 12.1/115 7.4/43.0 
Connectors com secção transversal redonda foram assumidas nestes cálculos. 
Veja uma pressupostos e condições no texto. 
Por outro lado, a maior resistência contra a trinca subcrítico crescimento do feldspathic verniz leva a menos rigorosas dimensões conector para a Inc - VM7 e TZP - CerS compósitos (Tabela 4). No caso de curto 3 - unidade pontes, a probabilidade de fratura do verniz sob condições subcrítico é marcadamente baixa (<5%) para o conector diâmetros iguais ou superiores a 4 milímetros. Para mais 4 - e 5 - unidade pontes, o conector dimensões exigidas para diminuir a probabilidade de insucesso dentro de 20 anos são mais elevados do que normalmente realizáveis na prática durante a preparação prótese. Portanto, é bem provável que rachaduras no verniz camada ocorrer em 4 e 5 - unidade póstero pontes dentro de um período de 20 anos. No entanto, é importante notar que a fratura do verniz camada não leva a concluir fracasso da ponte no caso de Inc - VM7 e TZP - CerS compósitos [23], devido a rachadura prender no verniz - quadro Interface. No entanto, rachaduras no verniz camada pode expor o quadro núcleo de um ambiente aquoso e, eventualmente, aumentar o crescimento subcrítico rachadura no quadro materiais. Por esse motivo, o crescimento da trinca subcrítico dentária quadro materiais sob condições molhadas e cíclico é abordada em publicações posteriores [19]. 
4. Conclusões 
A análise, in vitro vigência do verniz - quadro compósitos revelou que a água - assistida subcríticos rachadura crescimento conduz à propagação da trinca verniz superficial em estresse níveis sensivelmente mais baixos do que o verniz inertes força. Apatite - baseado folheados (Eris 2) foram observados para ser mais susceptíveis ao crescimento subcrítico rachadura que feldspathic folheados (CerS e VM7), provavelmente devido ao seu elevado dissolução na água sob tensão carregamento condições. 
Com base nestes resultados, in vitro, Emp2 - Eris compósitos não são adequados para materiais altamente carregado dentária pontes na região molares devido à elevada susceptibilidade à trinca subcrítico crescimento da apatite verniz, em associação com a ausência de detenção rachadura no verniz - Interface quadro. O feldspathic folheados da TZP - CerS e Inc - VM7 compósitos, por outro lado, pode apresentar vida útil superior a 20 anos se o conector de 3 - unidade pontes é adequadamente concebidos. No caso de 4 - e 5 - unidade póstero pontes, o conector dimensões exigidas para prorrogado vidas são difíceis de alcançar na prática, o que pode levar a pequenas fissuras na camada verniz devido ao crescimento subcrítico rachadura. However, the veneer cracks in the gingival side of the bridge connector do not result in complete failure of the TZP-CerS and Inc-VM7 composites, due to their ability to stop cracks at the veneer–framework interface. Therefore, the fracture of the veneer only exposes the core framework to the aqueous environment of the mouth. The long-term lifetime of zirconia-based bridges (TZP-CerS and Inc-VM7) is then governed by the water-assisted subcritical crack growth behavior of the framework material, which is the subject of a following publication. 

Acknowledgements 
The authors would like to gratefully acknowledge Madeleine Schumacher and Olivier Loeffel (University of Zürich, Center for Dental & Oral Medicine) for the kindness and valuable guidance during preparation of the specimens, as well as Marina Tambasco Studart for the patience and precision while layering most of the specimens evaluated in this study. We are also very thankful to Degudent, Vita Zahnfabrik and Ivoclar Vivadent for kindly supplying their framework and veneering materials for the TZP-CerS, Inc-VM7 and Emp2-Eris composites, respectively. 

References 
[1] F. Filser, H. Lüthy, P. Scharer and L. Gauckler, All-ceramic restorations by new direct ceramic machining process (DCM), J Dent Res 77 (1998), p. 762 [Abstract No. 1046]. 
[2] J.R. Kelly, Ceramics in restorative and prosthetic dentistry, Annu Rev Mater Sci 27 (1997), pp. 443–468. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (83) 
[3] H. Lüthy, F. Filser, L. Gauckler and P. Scharer, High reliable zirconia bridges by direct ceramic machining process (DCM), J Dent Res 77 (1998), p. 762 [Abstract No. 1045]. 
[4] B.R. Lawn, A. Pajares, Y. Zhang, Y. Deng, M.A. Polack and I.K. Lloyd et al., Materials design in the performance of all-ceramic crowns, Biomaterials 25 (2004), pp. 2885–2892. SummaryPlus | Full Text + Links | PDF (275 K) | View Record in Scopus | Cited By in Scopus (27) 
[5] W.S. Oh and K.J. Anusavice, Effect of connector design on the fracture resistance of all-ceramic fixed partial dentures, J Prosthet Dent 87 (2002), pp. 536–542. Abstract | PDF (148 K) | View Record in Scopus | Cited By in Scopus (17) 
[6] W. Oh, N. Gotzen and K.J. Anusavice, Influence of connector design on fracture probability of ceramic fixed-partial dentures, J Dent Res 81 (2002), pp. 623–627. View Record in Scopus | Cited By in Scopus (19) 
[7] J.R. Kelly, J.A. Tesk and J.A. Sorensen, Failure of all-ceramic fixed partial dentures in-vitro and in-vivo—analysis and modeling, J Dent Res 74 (1995), pp. 1253–1258. View Record in Scopus | Cited By in Scopus (93) 
[8] F. Filser, P. Kocher, F. Weibel, H. Lüthy, P. Scharer and L.J. Gauckler, Reliability and strength of all-ceramic dental restorations fabricated by direct ceramic machining (DCM), Int J Comput Dent 4 (2001), pp. 89–106. View Record in Scopus | Cited By in Scopus (42) 
[9] N. Goetzen, G. Natt, M.M. Morlock and J. Tinschert, Numerical stress and reliability analysis of all-ceramic fixed-partial dentures 2003 Summer Bioengineering Conference, June 25–29, Key Biscayne, Florida (2003), pp. 869–870. 
[10] S.O. Koutayas, M. Kern, F. Ferraresso and J.R. Strub, Influence of framework design on fracture strength of mandibular anterior all-ceramic resin-bonded fixed partial dentures, Int J Prosthodont 15 (2002), pp. 223–229. View Record in Scopus | Cited By in Scopus (7) 
[11] M. Kern, W.H. Douglas, T. Fechtig, J.R. Strub and R. Delong, Fracture strength of all-porcelain, resin-bonded bridges after testing in an artificial oral environment, J Dent 21 (1993), pp. 117–121. Abstract | View Record in Scopus | Cited By in Scopus (26) 
[12] A. Sundh, M. Molin and G. Sjogren, Fracture resistance of yttrium oxide partially-stabilized zirconia all-ceramic bridges after veneering and mechanical fatigue testing, Dent Mater 21 (2005), pp. 476–482. SummaryPlus | Full Text + Links | PDF (174 K) | View Record in Scopus | Cited By in Scopus (16) 
[13] H. Lüthy, F. Filser, O. Loeffel, M. Schumacher, L.J. Gauckler and C.H.F. Hammerle, Strength and reliability of four-unit all-ceramic posterior bridges, Dent Mater 21 (2005), pp. 930–937. View Record in Scopus | Cited By in Scopus (13) 
[14] T.A. Michalske and S.W. Freiman, A molecular mechanism for stress-corrosion in vitreous silica, J Am Ceram Soc 66 (1983), pp. 284–288. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (125) 
[15] J. Chevalier, C. Olagnon and G. Fantozzi, Subcritical crack propagation in 3Y-TZP ceramics: static and cyclic fatigue, J Am Ceram Soc 82 (1999), pp. 3129–3138. View Record in Scopus | Cited By in Scopus (33) 
[16] D.S. Jacobs and I.W. Chen, Cyclic fatigue in ceramics—a balance between crack shielding accumulation and degradation, J Am Ceram Soc 78 (1995), pp. 513–520. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (36) 
[17] R.H. Dauskardt, W. Yu and R.O. Ritchie, Fatigue crack-propagation in transformation-toughened zirconia ceramic, J Am Ceram Soc 70 (1987), pp. C248–C252. 
[18] T. Fett, D. Munz and G. Thun, Influence of frequency on cyclic fatigue of coarse-grained Al2o3, J Mater Sci Lett 12 (1993), pp. 220–222. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (7) 
[19] A.R. Studart, F. Filser, P. Kocher and L.J. Gauckler, Fatigue of zirconia under cyclic loading in water and its implications for the design of dental bridges, Dent Mater 23 (2007), pp. 106–114. SummaryPlus | Full Text + Links | PDF (545 K) | View Record in Scopus | Cited By in Scopus (4) 
[20] H. Fischer, M. Weber and R. Marx, Lifetime prediction of all-ceramic bridges by computational methods, J Dent Res 82 (2003), pp. 238–242. View Record in Scopus | Cited By in Scopus (20) 
[21] U. Lohbauer, A. Petschelt and P. Greil, Lifetime prediction of CAD/CAM dental ceramics, J Biomed Mater Res 63 (2002), pp. 780–785. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (13) 
[22] J. Geis-Gerstorfer, P. Fassler and R. Kirmeier, Fatigue behavior of three all ceramic materials, J Dent Res 81 (2002), p. A469 [Abstract No. 3835]. 
[23] A.R. Studart, F. Filser, P. Kocher, H. Lüthy and L.J. Gauckler, Mechanical and fracture behavior of veneer–framework composites for all-ceramic dental bridges, Dent Mater 23 (2007), pp. 115–123. SummaryPlus | Full Text + Links | PDF (577 K) | View Record in Scopus | Cited By in Scopus (2) 
[24] A.G. Evans and E.R. Fuller, Crack-propagation in ceramic materials under cyclic loading conditions, Metall Trans 5 (1974), pp. 27–33. View Record in Scopus | Cited By in Scopus (98) 
[25] D. Munz and T. Fett, Ceramics: mechanical properties, failure behaviour, materials selection, Springer-Verlag, Berlin (1999). 
[26] V. Imbeni, J.J. Kruzic, G.W. Marshall, S.J. Marshall and R.O. Ritchie, The dentin–enamel junction and the fracture of human teeth, Nat Mater 4 (2005), pp. 229–232. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (20) 
[27] P. Vanlanduyt, F. Li, J.P. Keustermans, J.M. Streydio, F. Delannay and E. Munting, The influence of high sintering temperatures on the mechanical-properties of hydroxylapatite, J Mater Sci-Mater Med 6 (1995), pp. 8–13. View Record in Scopus | Cited By in Scopus (72) 
[28] M. Guazzatto, K. Proos, L. Quach and M.V. Swain, Strength, reliability and mode of fracture of bilayered porcelain/zirconia (Y-TZP) dental ceramics, Biomaterials 25 (2004), pp. 5045–5052. 
[29] R. Schmitt, T. Fett and D. Munz, Cyclic fatigue of zirconia, Fatigue Fract Eng Mater Struct 19 (1996), pp. 1411–1420. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (7) 
[30] T. Fett, G. Martin, D. Munz and G. Thun, Determination of Da/Dn-Delta-K1 curves for small cracks in alumina in alternating bending tests, J Mater Sci 26 (1991), pp. 3320–3328. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (19) 
[31] S. Raynaud, E. Champion, D. Bernache-Assolant and D. Tetard, Dynamic fatigue and degradation in solution of hydroxyapatite ceramics, J Mater Sci-Mater Med 9 (1998), pp. 221–227. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (18) 
[32] S.M. Wiederhorn and L.H. Bolz, Stress corrosion and static fatigue of glass, J Am Ceram Soc 53 (1970), p. 543. Full Text via CrossRef 
[33] S.V. Dorozhkin, A review on the dissolution models of calcium apatites, Prog Cryst Growth Charact Mater 44 (2002), pp. 45–61. Abstract | Abstract + References | PDF (1361 K) | View Record in Scopus | Cited By in Scopus (19) 
[34] Y.M. Chen and X.G. Miao, Effect of fluorine addition on the corrosion resistance of hydroxyapatite ceramics, Ceram Int 30 (2004), pp. 1961–1965. SummaryPlus | Full Text + Links | PDF (360 K) | View Record in Scopus | Cited By in Scopus (7) 
[35] N.H. de Leeuw, Resisting the onset of hydroxyapatite dissolution through the incorporation of fluoride, J Phys Chem B 108 (2004), pp. 1809–1811. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (11) 
[36] J.F. Esquivel-Upshaw, K.J. Anusavice, H. Young, J. Jones and C. Gibbs, Clinical performance of a lithia disilicate-based core ceramic for three-unit posterior FPDs, Int J Prosthodont 17 (2004), pp. 469–475. View Record in Scopus | Cited By in Scopus (12) 
[37] M. Beghini, L. Bertini and F. Frendo, Measurement of coatings’ elastic properties by mechanical methods: part 1. Consideration on experimental errors, Exp Mech 41 (2001), pp. 293–304. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (4) 
[38] S.P. Timoshenko and J.M. Gere, Mechanics of materials, Litton Educational Publishing, Inc., London (1972). 

Appendix A. 
The stress on the veneer surface of multi-unit bridges at the gingival side of the connector was estimated using beam theory calculations assuming free-rotating supported beams. A cylindrical beam displaying the same diameter as a round bridge connector was taken as a representative model. Fig. A.1 schematically shows the stress and strain profiles along the z direction (y = 0) for the model composite beams considered in the calculations. The centroid (neutral axis) of the composite beam coincides in this case with the specimen geometric center, leading to a symmetric (but oppositely signed) stress and strain distribution (Fig. A.1). 
  
Display Full Size version of this image (29K) 

Fig. A.1. Stress (σ) and strain ( ) distribution along the z direction (y = 0) for a model cylindrical beam consisting of a framework core (in white) and an outer veneer layer (in grey). In this case, the bending moment M is applied in the xz plane. 
Based on beam theory, the longitudinal strain ( x) along the model composite beam for an applied moment M in the xz plane is given by [37] and [38]: 

  
(A.1) 
where Ek and Jk are the elastic modulus and the moment of inertia of the kth component of the composite beam, respectively (in this case, the veneer layer and the framework core). Taking into account that the elastic modulus of the framework (Eframe) is considerably higher than that of the veneer material (Eveneer), the contribution of the veneer coating to the overall stiffness ( ) was neglected in the calculations, leading us to the following simplified form of Eq. (A.1): 

  
(A.2) 
The strain at the veneer–framework interface ( int) is then given by Eq. (A.2) for z = rframe, whereas the strain on the veneer surface ( surf) is obtained when z = (rframe + hveneer). From these strain values, the stress along the composite beam can be calculated using Hooke's law (σ = E ): 

  
(A.3) 

  
(A.4) 

  
(A.5) 
where , and are the stress values indicated in Fig. A.1. 
By substituting the appropriate strain values from Eq. (A.2) into Eqs. (A.3) and (A.5), and subsequently dividing Eq. (A.5) by Eq. (A.3), the following relation can be obtained for the stress on the veneer surface: 

  
(A.6) 
where rframe and hveneer are the radius of the framework core and the thickness of the veneer layer, respectively (Fig. A.1). The stress on the veneer surface was estimated from this equation, based on the maximum stress values calculated on a previous study [19] for the framework core ( ). 


  Correspondence to: Department of Materials, Nonmetallic Inorganic Materials, Wolfgang–Pauli-Strasse 10, ETH Hönggerberg, HCI G 539, 8093 Zürich, Switzerland. Tel.: +41 44 632 37 18; fax: +41 44 632 11 32. 



Dental Materials 
Volume 23, Issue 2, February 2007, Pages 177-185