Light transmission through all-ceramic dental materials: a pilot study

Flavio H Rasetto BDS, MSa, , , Carl F Driscoll DMDb, Vince Prestipino DDSc, Radi Masri BDS, MSd and Joseph Anthony von Fraunhofer MS, PhDe 
aAssistant Professor, Advanced Education Program in Prosthodontics, Department of Restorative Dentistry USA
bAssociate Professor, Director, Advanced Education Program in Prosthodontics, Department of Restorative Dentistry USA
cPrivate Practice, Bethesda, Md USA
dClinical Instructor, Advanced Education Program in Prosthodontics, and PhD candidate USA
eProfessor, Director of Biomaterials, Department of Restorative Dentistry USA 
Baltimore College of Dental Surgery, University of Maryland, Baltimore, Md.  Available online 19 May 2004.

- Abstract

- Statement of problem

Light transmission through veneers affects the degree of polymerization of light-polymerized resin luting agents.

- Purpose

Light transmission through different all-ceramic veneers when irradiated by 3 types of light-polymerization units was measured to evaluate the degree of polymerization.

- Material and methods

The power outputs from a conventional halogen (3M Unitek), a plasma arc (Apollo 95E), and a high-intensity halogen (Kreativ Kuring Light Model 2000) light were measured by a radiometer. The light intensity (mW/cm2) from these units was also measured after transmission through 0.25-, 0.40-, and 0.60-mm–thick Procera copings and through 1-mm-thick disks of feldspathic porcelain (Ceramco II), aluminous porcelain (Vitadur Alpha), and a castable pressed ceramic (IPS Empress). Two disks of each material were fabricated by a commercial laboratory in accordance with the manufacturers' recommendations. The light transmissions through these materials were compared with those through Procera–aluminous porcelain combinations. These disks were fabricated by applying appropriate thicknesses of aluminous porcelain (AllCeram). Light intensities were subjected to ANOVA and post hoc Scheffé tests at a priori α=.05.

- Results

Intensities of light from 3 polymerization units, conventional halogen light, high-intensity halogen light, and plasma arc, were 660, 1050, and 2475 mW/cm,2 respectively, and these together with the ceramic veneer thickness dictated the light transmission through veneers. Mean values (±SD) of light transmission through Procera copings for thicknesses of 0.25, 0.40 and 0.60 mm, respectively, were as follows: Apollo plasma arc (1083 ± 117, 843 ± 59, and 593 ± 132); Kreativ high-intensity halogen (425 ± 41, 345 ± 74, and 256 ± 79); 3M Unitek halogen (270 ± 102, 230 ± 48, and 180 ± 113). Comparable transmissions occurred with the other ceramics. Only the plasma arc and the high-intensity halogen polymerization units emitted light of sufficient energy to effect polymerization of a resin luting agent (P=.045).

- Conclusion

Within the limitations of this study, the intensity of light transmitted through ceramic veneers was dictated by the polymerization unit and the type and thickness of the ceramic. With conventional halogen polymerization units, there may be insufficient light transmission through thicker veneers or all-ceramic crowns for adequate light polymerization.

- Article Outline

1. Material and methods

2. Results

3. Discussion

4. Conclusions

References

Clinical implications

This in vitro pilot study demonstrated that polymerization of light-polymerized resin luting agents beneath ceramic restorations may be incomplete when lower-output polymerization units are used.

Esthetic demands of patients and clinicians have led to the development of all-ceramic crowns. Bonded all-ceramic restorations provide superior esthetics because ceramics allow diffuse transmission as well as diffuse and specular reflectance of light, reproducing a depth of translucency and color mimicking that of natural teeth.1

There are several commercially available all-ceramic systems that are used for the fabrication of complete coverage crowns as well as veneers. Feldspathic porcelain is used primarily for metal-ceramic restorations. When used as a veneer material, etched and bonded to enamel, it is a predictable restoration.2 Feldspathic porcelain generally does not have adequate strength to be used for all-ceramic complete coverage restoration.

Aluminous porcelain was developed by McLean and Hughes3 in 1965, and this material, which utilized alumina dispersions within the vitreous matrix, was the first clinically acceptable approach to achieving an esthetic high-strength dental ceramic. In the early 1990s, a pressable ceramic system was developed by Wohlwend et al4 that combined the lost-wax casting technique for metals with modern ceramic technology. In this technique, the restoration wax pattern is invested and burned out in a furnace and a special ceramic material is then “cast” under pressure into the mold. The ceramic has a high leucite content core, such that the pressed or cast restoration has the requisite strength for both the anterior and posterior restorations. The leucite core may be contoured to permit veneering, commonly with the same ceramic material, or it can be stained to achieve optimal esthetics.

Another high-strength esthetic all-ceramic system is the Procera AllCeram crown (Procera AllCeram; Nobel Biocare, Yorba Linda, Calif), introduced in 1993 by Andersson and Oden.5 In the Procera system, the core material consists of a densely sintered high-purity alumina (99.9%) that can be veneered with porcelain to achieve optimal esthetics. The coping fabrication process utilizes CAD-CAM technology and a scanned image of the original master die to produce a larger milled die, compensating for 20% shrinkage of the coping that occurs during presintering at high temperature. Consequently, the coping strength is reported to be comparable among all-ceramic restorations.6 An additional advantage of Procera AllCeram copings is that their radiopacity is similar to that of dentin. This characteristic makes it possible to diagnose changes in the underlying tooth structure supporting the coping.

Although coping strength is an important characteristic, clinical success with bonded restorations and veneers is multifactorial.7 The variables that contribute to a durable and predictable ceramic veneer restoration include the adhesive material, its polymerization behavior and mechanism, the light polymerizing unit, and the nature and thickness of the ceramic veneer material itself.

Resin luting agents are required for luting of ceramic crowns to assure high bond strengths,8 and their translucency and minimal film thickness provide optimum esthetics and enhance the clinical survival rate of all-ceramic restorations.9 It should be noted, however, that the alumina core of Procera crowns permits the use of both resin luting agents and conventional cements, for example, zinc phosphate, for cementation.10 The type of resin luting agent used for luting all-ceramic crowns affects the clinical performance of the restoration.11 Light-polymerized resins are often preferred to chemical- and dual-polymerized materials because of their color stability and their “polymerization-on-demand” characteristics.12 However, adequate polymerization of resin luting agents is essential to ensure that the adhesive achieves proper strength and provides an acceptable marginal seal, thereby preventing microleakage.11 It is accepted that the degree of double bond conversion is 50% to 70% with light and chemically polymerizing systems.13. and 14. It is clear that any reduction in the degree of conversion due to inadequate power output from a polymerization lamp will have an adverse effect on resin performance. The thickness and shade of the all-ceramic restoration affects the degree of polymerization of the light-polymerized resin luting agent. Generally, the thicker the restoration or the darker its shade, the more critical the intensity of the incident light to achieve optimal photopolymerization of the material.12 In particular, the amount of light transmission through the ceramic restoration affects the potential for polymerization of the resin luting agent.15., 16. and 17. A variety of light polymerization units are available, and in 1993, some of these lamps were evaluated for their power output and emitted light wavelength.18 It was found that the output power and wavelength of 5 conventional (halogen) lights were very similar over the wavelength band (450-500 n·m) customarily used for light polymerization. The authors evaluated the light units spectro-radiometrically, with light intensities being determined at specific wavelengths, and the results were compared with the instantaneous readings of power density measured with a commercial radiometer. The data from the 2 techniques correlated well over the wavelength range 450 to 500 n·m, and it was concluded that a radiometer was an effective tool to rapidly characterize the polymerization efficiency of dental light units.18 The authors indicated that the power outputs of the 5 commercial dental polymerization lights were 100, 250, 340, 390, and 580 mW/cm2.18 It should be noted, however, that a decrease in line voltage has been shown to adversely affect the intensity of polymerization light units,19. and 20. and this could be a factor in studies of polymerization light outputs.

Because it has been suggested that 1-mm composite increments require irradiation for 60 seconds at a light output power of at least 400 mW/cm2 to achieve adequate polymerization,21 only 2 of the 5 units tested by Lee et al18 appear to be satisfactory for clinical use. The ISO-recommended intensity for polymerization lights is 300 mW/cm2, and the standard depth-of-polymerization requirement is 1.5 mm.22 It was stated that “polymerization lights with an intensity of 300 mW/cm2 appear to effectively polymerize most resin-based composite materials when appropriate polymerization times are used, which, in some cases, are longer than those recommended by the manufacturers.”22 In fact, a recent study23 has indicated that irradiation times used by dentists for light-polymerizing cements are too short. Longer polymerization times are necessary to offset decreases in light intensity incident upon the resin adhesive due to both the overlying ceramic material and light source factors in order to achieve an adequate degree of polymerization. Other authors24 have suggested that whereas porcelain opacity did not significantly affect the hardness of resin luting agents polymerized through porcelain veneers, dual polymerization materials and prolonged polymerization times might be essential for clinical success.

Recently, dental polymerization units with higher outputs have been developed to reduce the time required to polymerize resin cements. Manufacturers of these units claim that adequate polymerization of a 2-mm-thick layer of resin may be achieved in 3 to 5 seconds, compared to the 40-second polymerization time required with a conventional light polymerization unit.25 The literature, however, tends to be conflicting with regard to these claims. Two studies26. and 27. showed that the degree of polymerization and bond strengths achieved when resins were polymerized for 6 to 9 seconds with a plasma arc light were comparable to that obtained with halogen polymerization for 40 seconds. It was suggested that the shade or opacity of porcelain had little effect on the bond strengths when the resin luting agent was polymerized through porcelain veneers. Other authors,28 however, have reported that the depth of polymerization was always greater with conventional (halogen light) polymerization and a “soft start” regimen. The latter is a procedure wherein the initial stage of the light polymerization occurs at low intensity and the intensity of the light output is then increased to maximum intensity to accelerate the rate of polymerization reaction.

Since both the type of ceramic and the veneer thickness are important parameters in light transmission, it is clear that the clinical performance of light-polymerized resins will be determined by both the polymerizing unit output and the amount of light that is transmitted through the restoration to the resin-based adhesive. The purpose of the present study was to measure light transmission through different thicknesses of copings fabricated from a sintered high-alumina ceramic (Procera). In addition, the translucencies of Procera AllCeram system veneers were compared to the translucencies of other all-ceramic systems.

- 1. Material and methods

In Phase I of this study, the light transmission through core materials of the Procera AllCeram System was evaluated using 0.25-, 0.40-, and 0.60-mm-thick Procera disks, 10 mm in diameter. In Phase II of the study, the light transmissions through core and veneering material of all-ceramic systems were compared using 1-mm disks fabricated from different all-ceramic systems.

For Phase I, a total of 6 disks, 2 disks of each thickness, were fabricated by the manufacturer (Nobel Biocare, Yorba Linda, Calif). The disks were milled to the desired thickness by the manufacturer. Three dental polymerization light systems were evaluated: a conventional halogen light (3M Unitek; 3M Inc, St. Paul, Minn), a plasma arc light (Apollo 95E; Dental Medical Diagnostic Systems Inc, Woodland Hills, Calif), and a high-intensity halogen light (Kreativ Kuring Light Model 2000; Kreativ Inc, San Diego, Calif). Each polymerization light had a light guide with 7-mm-diameter tips. The intensities of the light outputs from each polymerization unit were assessed with the same radiometer prior to each use in the present study. The built-in radiometer from the Kreativ high-polymerization light unit was tested by using a second radiometer (Cure Lite; EFOS Inc, Williamsville, NY) to confirm accuracy.

Light transmission by the veneers was determined by placing each in turn over the internal radiometer of the Kreativ light unit and irradiating the specimen for 10 seconds at the highest setting of each light polymerization (Fig. 1). Over the irradiation period of 10 seconds, the highest and lowest values of the transmitted light were recorded. The instrument measured the transmitted light through each veneer as the luminous flux in mW/cm2 received by the radiometer detector. The average of these 2 readings was used to compare the light transmission for each disk thickness and polymerization light combination. The intensities of transmitted light were compared by 1-way ANOVA with post hoc Scheffé tests (α=.05).

|  | Display Full Size version of this image (6K) |
| --- | --- |

Fig. 1. Radiometer measurement of light polymerization through ceramic disk.

After the Phase I study was completed, the Apollo lamp used in that portion of the study had to be returned to the manufacturer; Phase II of the study was continued with a different Apollo light unit of the same model. It was noted, however, that the light outputs for the 2 lamps differed (P<.05) and were, respectively, 2450 to 2500 mW/cm2 and 1550 to 1650 mW/cm2 for the Phase I and II lamps. Thus, the luminous flux for the Phase II lamp was 34% lower than the output light intensity achieved with the Phase I lamp.

For Phase II of the study, specimen disks 1.0-mm-thick × 10-mm-diameter were fabricated from feldspathic porcelain (Ceramco II; Dentsply Ceramco, Burlington, NJ), aluminous porcelain (Vitadur Alpha; Vident Inc, Brea, Calif), and a castable pressed ceramic (IPS Empress; Ivoclar Vivadent, Amherst, NY). Two disks of each material were fabricated by a commercial laboratory in accordance with the manufacturers' recommendations. The light transmissions through these materials were compared with those through Procera-aluminous porcelain combinations. These disks were fabricated by applying appropriate thicknesses of an aluminous porcelain (Procera AllCeram; Nobel Biocare) onto the 0.25- and 0.60-mm-thick Procera disks used in Phase I, to produce disks with a total thickness of 1.0 mm. These combination disks were fabricated according to the manufacturers' recommendations. The thickness of each disk was measured by an electronic caliper (Mitutoyo, Tokyo, Japan) to ±0.1 mm to ensure standardized specimen thicknesses of 1 mm.

The 6 disks for Phase II were divided into 3 groups of 2 disks each, and these were tested following the same protocol utilized in Phase I of the study. As this was a pilot study, few specimens were available for measurement, and no power analysis to determine adequate sample size was performed. Accordingly, the data were subjected to a 1-way ANOVA and post hoc Scheffé test at (α=.05) to determine, and identify, any differences within all collected data.

- 2. Results

Radiometer measurements of the transmitted light from photopolymerizing light units after passage through the 3 thicknesses of Procera copings are presented in Table I. It should be noted that the recommended minimum light intensity required for activation of camphoroquinone is 280 to 300 mW/cm2.22

Table I.

Radiometer measurements (mean ± SD) of polymerization light intensity after passing through Procera copings (mW/cm2) (n=2)

|  | Apollo | Kreativ | 3M Unitek |
| --- | --- | --- | --- |
| Output power (mW/cm2) | 2450-2500 | 1040-1080 | 610-710 |
| Coping thickness (mm) |  |  |  |
| 0.25 | 1083 ± 117 | 425 ± 41 | 270 ± 102 |
| 0.40 | 843 ± 59 | 345 ± 74 | 230 ± 48 |
| 0.60 | 593 ± 132 | 256 ± 79 | 180 ± 113 |

Light transmission through the Procera copings was reduced with increased coping thicknesses for all 3 lights. The decrease in transmission was 14 to 22% for the 0.40-mm coping (Apollo 843 ± 59, Kreativ 345 ± 74, 3M 230 ± 48) compared to the 0.25-mm copings (Apollo 1083 ± 117, Kreativ 425 ± 41, 3M 270 ± 102). There was an additional drop in light transmission of 22% to 30% between the 0.40- and the 0.60-mm copings (Apollo 593 ± 132, Kreativ 256 ± 79, 3M 180 ± 113). The decreased intensity of the light transmitted through the 3 copings was significant (P=.045) for the Apollo light. However, no significant differences (P=.180) were found for the transmitted light through the 3 thicknesses of copings for the Kreativ and 3M (conventional halogen) polymerization units.

When light transmission for the 3 lights through the same thickness of coping was compared, it was found that the Apollo light produced a significantly (P<.05) greater transillumination than the other 2 lights for all 3 coping thicknesses. No differences (P=.162) were found in the light transmission for the Kreativ and 3M lights at all 3 coping thicknesses.

The intensities of transmitted light for the 3 polymerization lights units through 1-mm-thick veneers of Procera, IPS Empress, Vita Alpha, and a feldspathic porcelain are given in Table II. The amount of light transmitted through the 2 thicknesses of Procera copings (0.25, 0.60, Table II) from the halogen lamp and Kreativ light unit through 0.60 Procera copings (Table II) was below the detection limit of the radiometer (100 mW/cm2). No statistically significant differences in light transmission through the 5 types of ceramic were found for any of the lights. Likewise, no differences (P=.180) were found in light transmission for the Kreativ and Halogen lamps for any of the ceramic materials. Although there were no differences (P>.05) in light transmission for all 5 ceramic materials for the Apollo and Kreativ lamps, the transmitted light from the Apollo light unit was greater (P<.05) than for the halogen light with all the test materials. Interestingly, the light transmission through the Vita Alpha and Feldspathic porcelain disks was identical for the Apollo light. Differences in transmission through the 2 materials might become more evident when larger sample sizes are used.

Table II.

Radiometer measurements (mean ± SD) of polymerization light intensity (mW/cm2) after passing through 1-mm-thick porcelain copings

| Porcelain discs (n=2) | Apollo (1550-1650) | Kreativ (1050-1060) | 3M Unitek (580-710) |
| --- | --- | --- | --- |
| 0.25 (Procera)* | 535 ± 50 | 175 ± 106 | — |
| 0.60 (Procera)* | 340 ± 62 | — | — |
| IPS Empress | 640 ± 65 | 355 ± 127 | 265 ± 99 |
| Vita Alpha | 590 ± 113 | 325 ± 40 | 250 ± 71 |
| Feldspathic | 590 ± 113 | 335 ± 50 | 230 ± 35 |

—Radiometer did not reach minimum of 100 mW/cm2.
* Thickness of Procera coping with porcelain added to reach total veneer thickness of 1 mm.

It was noted during the second phase of the study, which of necessity had to be performed at a different location than that of Phase I, that the intensity of emitted light from the Apollo 95E light used in the new location decreased by 34% (1550-1650 mW/cm2 compared to 2450-2500 mW/cm2, Table I and Table II). This change in emitted light was ascribed to differences in the mains or line electrical voltage for the 2 test locations (The Dental School, University of Maryland, Baltimore, and a private dental office in Bethesda, Md). The Kreativ and 3M lights were tested in one location only, and no comments can be made regarding changes in light intensity output with variation in mains voltage.

- 3. Discussion

Bonded all-ceramic restorations are popular with clinicians and patients because of their superior esthetics: notably the diffuse transmission combined with diffuse and specular reflectance of light that mimics the depth of translucency and color of natural teeth. As a result, the ceramic veneer has become popular for improving esthetics. However, clinical success with these restorations requires that they are bonded to teeth with high-strength, low-solubility, and thin-film adhesives, commonly light-polymerized resin luting agents.8 Implicit in this is the requirement that the resin luting agent is completely polymerized within an acceptable period of time.19. and 21.

Clinicians commonly evaluate the degree of polymerization of resin luting agents by probing with an explorer. The resin luting agent is assumed to be “set” when it exhibits satisfactory resistance to probing, although the material may not, in fact, have been completely polymerized. The incomplete light polymerization of resin luting agents can arise because of a number of clinical and technical parameters that limit the efficacy of light polymerizing systems. In particular, the restoration thickness varies at different regions; for example, the veneer thickness may be greater than 1 mm near the incisal edge but only 0.3 mm in the gingival one third. Consequently, although the tip of the light wand may be in direct contact with the veneer surface, the distance between the tip and the surface of the resin luting agent can vary markedly over the entire veneer. Thus, the total irradiance provided by the light wand would be less in the incisal region than at locations where the ceramic veneer is thinner, and this may affect the light polymerization of the resin luting agent. Further, differences in veneer translucency arising from compositional or processing variables may also affect the light polymerization reaction.

In addition, for ceramic veneers of comparable thickness and composition, the light polymerization of the underlying resin luting agent will be markedly affected by the light source, notably the wavelength and intensity of emitted light. While the wavelength of the light is determined by the generating source (the bulb or plasma arc, for example), factors such as the age of the light source, the characteristics of the guiding wand, and the separation between wand tip and the veneer all affect the intensity of emitted light and the efficiency of light polymerization. Another factor can often be the mains voltage, which is used to power the light source. Although the mains voltage is nominally 220V, customary fluctuations are in the range of ± 10% and can be as great as 20%19; that is, the lamp may be required to provide an adequate output at operating voltages as low as 180V and as high as 260V. Irradiating a resin luting agent through a medium-to-thick veneer using a low-intensity light may not provide sufficient light transmission to effect initiation of the camphoroquinone (CQ). This was shown by the findings for the conventional tungsten/halogen and Kreativ lights with the 0.60-mm Procera specimens, for which the resin luting agent did not achieve adequate surface hardness and, presumably, acceptable bond strength. In contrast, all specimens irradiated with the higher power Apollo 95E plasma arc light transmitted light of sufficient intensity to activate the CQ. Acceptable polymerization was found with this light, regardless of the fluctuations in the voltage of the mains electrical supply. In contrast, variations in the degree of polymerization were noted for the other polymerization lights, demonstrating the importance of the light output power.

When the light output of the Kreativ 2000 xenon/halogen high intensity light was 1050 to 1060 mW/cm2, sufficient light was transmitted through the 0.25- and 0.40-mm-thick Procera copings to activate CQ. However, with the 0.60-mm-thick Procera copings, there was insufficient light transmission with the Kreativ light to activate CQ. The conventional halogen light was incapable of providing sufficient light transmission through all 3 thicknesses of the Procera coping to effect satisfactory polymerization of the resin luting agent. It should be noted that the data reported here represent readings obtained from a single light unit of each type and only 2 specimens of each ceramic material. Nevertheless, the present findings suggest that adequate polymerization of resin luting agents through Procera copings and veneers may not be possible with conventional halogen polymerization lights unless longer irradiation times are used or the light unit has a high-efficiency bulb. The same reservations would appear to be justified with conventional halogen light sources when polymerizing resin luting agents through other coping materials, as noted from the data in Table II. Other researchers have reported similar findings, notably the need for longer polymerization times23. and 24. to ensure clinical success with conventional lights. These authors also suggested that dual polymerization adhesives might be advantageous. In addition, it has been reported26. and 27. that the polymerization achieved within 6 seconds with plasma arc and xenon lamps was comparable to a 40-second polymerization with a conventional halogen light, findings that support the data presented here. While it was observed26 that the strength of the veneer bond was relatively unaffected by the shade or opacity of the porcelain, longer polymerization times were required for thicker, more darkly shaded porcelains.

While these findings support those of the present study, it should be noted that other authors28 have suggested that greater depths of polymerization were obtained with conventional polymerization and soft-start polymerization than were achieved with plasma arc polymerization. Future studies should be undertaken with an increased number of specimens to enable a more comprehensive statistical analysis to be performed. The study should also be expanded to include other all-ceramic systems and dental light polymerization units, particularly the newly introduced LED polymerization units. The data generated by such studies should provide adequate information to facilitate understanding of the factors affecting light transmission through all-ceramic materials and their influence on subsequent clinical behavior.

- 4. Conclusions

The data from this in vitro pilot study indicate that the effective polymerization of an adhesive resin luting agent may be dictated by the type of polymerizing light, the nature of the ceramic veneer, and, notably, the thickness of the ceramic veneer. Based on the small sample size evaluated, there does not appear to be adequate transmission of light through thicker copings to achieve effective resin polymerization when the resin is light polymerized with conventional and lower-output high-intensity polymerization lights. Furthermore, light polymerization of resin luting agents through even thin Procera copings does not appear to be achievable with conventional halogen light units.

- References

1. B Touati and P Miara, Light transmission in bonded ceramic restorations, J Esthet Dent 5 (1993), pp. 11–18. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (14)

2. M Friedman, A 15-year review of porcelain veneer failure—a clinician's observations, Compend Cont Educ Dent 19 (1998), pp. 625–632.

3. J.W McLean and T.H Hughes, The reinforcement of dental porcelain with ceramic oxides, Br Dent J 119 (1965), pp. 251–267.

4. A Wohlwend, J.R Strub and P Scharer, Metal ceramic and all-porcelain restorations: current considerations, Int J Prosthodont 2 (1989), pp. 13–26. View Record in Scopus | Cited By in Scopus (17)

5. M Andersson and A Oden, A new all ceramic crown. A dense-sintered, high- purity alumina coping with porcelain, Acta Odontol Scand 51 (1993), pp. 59–64. View Record in Scopus | Cited By in Scopus (90)

6. W.C Wagner and T.M Chu, Biaxial flexure strength and indentation fracture toughness of three new dental core ceramics, J Prosthet Dent 76 (1996), pp. 140–144. SummaryPlus | Full Text + Links | PDF (1212 K) | View Record in Scopus | Cited By in Scopus (69)

7. W.D Cook, Spectral distributions of dental photopolymerization sources, J Dent Res (1982), pp. 1436–1438. View Record in Scopus | Cited By in Scopus (51)

8. S.F Rosenstiel, M.F Land and B.J Crispin, Dental luting agents: a review of the current literature, J Prosthet Dent 80 (1998), pp. 280–301. SummaryPlus | Full Text + Links | PDF (381 K) | View Record in Scopus | Cited By in Scopus (77)

9. K.A Malament and S.S Socransky, Survival of Dicor glass-ceramic dental restorations over 16 years Part III: effect of luting agent and tooth or tooth-substitute core structure , J Prosthet Dent 86 (2001), pp. 511–519. Abstract | PDF (195 K) | View Record in Scopus | Cited By in Scopus (31)

10. A Oden, M Andersson, I Krystek-Ondracek and D Magnusson, Five year clinical evaluation of Procera All Ceram crowns, J Prosthet Dent 80 (1998), pp. 450–456. SummaryPlus | Full Text + Links | PDF (157 K) | View Record in Scopus | Cited By in Scopus (87)

11. E Asmussen, Factors affecting the quantity of remaining double bonds in restorative resin polymers, Scand J Dent Res 90 (1982), pp. 490–496. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (80)

12. M.L Myers, W.F Caughman and F.A Rueggeberg, Effect of restoration composition, shade, and thickness on the cure of a photoactivated resin cement, J Prosthodont 3 (1994), pp. 149–157. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (20)

13. K Yoshida and E.H Greener, Effect of photoinitiator on degree of conversion of unfilled light-cured resin, J Dent 22 (1994), pp. 296–299. Abstract | View Record in Scopus | Cited By in Scopus (38)

14. K.J Anusavice and R.W Phillips, Phillips' science of dental materials (11th ed.), Saunders Elsevier Science, St Louis (2003) p. 415.

15. K.L O'Keefe, P.L Pease and H.K Herrin, Variables affecting the spectral transmittance of light through porcelain veneers, J Prosthet Dent 66 (1991), pp. 434–438. Abstract | View Record in Scopus | Cited By in Scopus (18)

16. F.A Rueggeberg and R.G Craig, Correlation of parameters used to estimate monomer conversion in a light-cured composite, J Dent Res 67 (1988), pp. 932–937. View Record in Scopus | Cited By in Scopus (127)

17. K.C Chan and D.B Boyer, Curing light-activated composite cement through porcelain, J Dent Res 68 (1989), pp. 476–480. View Record in Scopus | Cited By in Scopus (24)

18. S.Y Lee, C.H Chiu, A Boghosian and E.H Greener, Radiometric and spectroradiometric comparison of power outputs of five visible light-curing units, J Dent 21 (1993), pp. 373–377. Abstract | View Record in Scopus | Cited By in Scopus (22)

19. I.E Ruyter and H Oysaed, Conversion in different depths of ultraviolet and visible light activated composite materials, Acta Odontol Scand 40 (1982), pp. 179–192. View Record in Scopus | Cited By in Scopus (100)

20. P.L Fan, W.T Wozniak, W.D Reyes and J.W Stanford, Irradiance of visible light-curing units and voltage variation effects, J Am Dent Assoc 115 (1987), pp. 442–445. View Record in Scopus | Cited By in Scopus (20)

21. F.A Rueggeberg and W.F Caughman, The influence of light exposure on polymerization of dual-cure resin cements, Oper Dent 18 (1993), pp. 48–55. View Record in Scopus | Cited By in Scopus (50)

22. P.L Fan, R.M Schumacher, K Azzolin, R Geary and F.C Eichmiller, Curing-light intensity and depth of cure of resin-based composites tested according to international standards, J Am Dent Assoc 133 (2002), pp. 429–434. View Record in Scopus | Cited By in Scopus (25)

23. C Strydom, Curing lights—the effects of clinical factors on intensity and polymerization, SADJ 57 (2002), pp. 181–186. View Record in Scopus | Cited By in Scopus (5)

24. J.J Linden, E.J Swift Jr., D.B Boyer and B.K Davis, Photo-activation of resin cements through porcelain veneers, J Dent Res 70 (1991), pp. 154–157. View Record in Scopus | Cited By in Scopus (14)

25. A Peutzfeldt, A Sahafi and E Asmussen, Characterization of resin composites polymerized with plasma arc curing units, Dent Mater J 16 (2000), pp. 330–336. SummaryPlus | Full Text + Links | PDF (88 K) | View Record in Scopus | Cited By in Scopus (96)

26. K Watanabe, E Ohnishi, T Kaneshima, A Mine and H Yatani, Porcelain veneer bonding to enamel with plasma-arc light resin curing, Dent Mater J 21 (2002), pp. 61–68. View Record in Scopus | Cited By in Scopus (11)

27. L.J Oesterle, S.M Newman and W.C Shellhart, Rapid curing of bonding composite with a xenon plasma arc light, Am J Orthod Dentofacial Orthop 119 (2001), pp. 610–616. Abstract | PDF (205 K) | View Record in Scopus | Cited By in Scopus (41)

28. H Jung, K.H Friedl, K.A Hiller, A Haller and G Schmalz, Curing efficiency of defferent polymerization methods through ceramic restorations, Clin Oral Investig 5 (2001), pp. 156–161. View Record in Scopus | Cited By in Scopus (13)

Reprint requests to: Dr. Flavio H. Rasetto Rm. 3 D 08 Department of Restorative Dentistry Baltimore College of Dental Surgery University of Maryland 666 West Baltimore Street Baltimore, MD 21202

The Journal of Prosthetic Dentistry Volume 91, Issue 5, May 2004, Pages 441-446