Precision of fit: The Procera AllCeram crown, , , , , , , ,

Kenneth B. May DDS, MSa, Melinda M. Russell DDS, MSa, Michael E. Razzoog DDS, MS, MPHb and Brien R. Lang DDS, MSc 
School of Dentistry, University of Michigan Ann Arbor, Mich 

Available online 4 November 2005.

- Abstract

Statement of problem. Strength, color stability, and precision of fit are requirements for all-ceramic restorations. The Procera AllCeram crown system, composed of a densely sintered high-purity alumina core combined with a low fusing surface porcelain, appears to satisfy most of these requirements. However, evaluation of marginal fit has not been reported. Purpose. This study measured the precision of fit of the Procera AllCeram crown fabricated with Procera CAD/CAM technology for the premolar and molar teeth fit to a die. Material and methods. Five ivorine maxillary first premolars and first molars were prepared for full-coverage crowns. Preparations were standardized with a convergence angle of 10 degrees, chamfer margins of 1.3 to 1.5 mm circumferentially, and occlusal reduction of 2.0 mm. AllCeram crowns were fabricated for the dies, and the fit of the crown to the die was determined by using a standardized procedure with a silicone impression material that served a dual role: (1) as a retrievable luting agent, and (2) to replicate the internal aspects of the crown. Laser videography was used to measure the gap dimension between the crowns and the dies at the marginal opening, the axial wall, the cusp tip, and the occlusal adaptation measurement locations. Mean gap dimensions and standard deviations (SDs) were calculated for marginal opening, internal adaptation, and precision of fit.

Results. Mean gap dimensions and standard deviations at the marginal opening for the premolar and molar crowns were 56.0 μm SD ± 21 and 63.0 μm SD ± 13 μm, respectively. The mean gap dimensions and SDs of the internal adaptation were 69.0 μm SD ± 17 μm for axial wall, 48.0 μm SD ± 12 μm for cusp tip, and 36.0 μm SD ± 7 μm for occlusal adaptation for the premolar crowns; and 49.0 μm SD ± 3 μm axial wall, 67.0 μm SD ± 21 μm cusp tip, and 74.0 μm SD ± 29 μm occlusal adaptation for molar crowns. Precision of fit and SDs for premolar and molar crowns were 52.0 μm SD ± 19 μm and 63.0 μm SD ± 20 μm, respectively. Mean marginal openings and precision of fit gap dimensions for the crown groups were not significantly different at the .05 level. However, gap dimensions that defined the internal adaptation at the measurement locations were different (P≤.05).

Conclusion. Mean gap dimensions for marginal openings, internal adaptation, and precision of fit for the crown groups were below 70 μm. These findings show that the crowns studied can be prescribed with confidence knowing that the precision of fit will consistently be less than 70 μm. (J Prosthet Dent 1998;80:394-404.)

- Article Outline

MATERIAL AND METHODS

RESULTS

DISCUSSION

CONCLUSIONS

Acknowledgements

References

The findings of this in vitro study should aid the dentist in selecting an all-ceramic crown system that has the marginal fidelity that is required for a favorable clinical prognosis. If the crowns used in this investigation are an accurate representation of the crown fit achievable in the dynamic environment of the oral cavity, then these crowns can be prescribed for the esthetic zone and the posterior dentition with confidence that an acceptable clinical fit will be consistently achieved.

The demand for esthetic dental restorations in all quadrants has made the all-ceramic restoration a popular restoration. All-ceramic restorations must satisfy clinical requirements of strength, precision of fit, and color stability if they are to be successful restorations.1, 2, 3 and 4 To achieve the appropriate strength in an all-ceramic crown, high-purity alumina was suggested by McLean and Hughes5 as a material to reinforce the feldspathic porcelain matrix. The alumina-reinforced dental porcelain has resulted in the development of several new restorative materials.6, 7 and 8 The Procera AllCeram crown system (Nobel Biocare, Göteborg, Sweden), composed of a densely sintered high-purity aluminous oxide coping combined with a low fusing surface porcelain, was introduced by Nobel Biocare AB and Sandvik Hard Materials AB.9 and 10 This system requires 3 steps for fabrication: (1) creation of an enlarged die from a working die; (2) sintering a high-purity aluminous oxide powder at 1550°C on the die to form a Procera coping, and (3) completing the crown form by the addition of a low fusing surface porcelain to the coping in developing the anatomy of a tooth (Fig. 1).

|  | Display Full Size version of this image (55K) |
| --- | --- |

Fig. 1. Procera AllCeram crown restoration.

Wagner and Chu11 reported substantial differences in flexural strength for the Procera coping material (687 MPa) when compared with In-Ceram (352 MPa) and Empress (134 MPa) all-ceramic restorative systems. The strength of the Procera coping material, combined with the low fusing surface porcelain, was tested in tension and compression with a biaxial flexural test (similar to ASTM F394-78 and ISO/IDS 6872 tests).12 The thickest porcelain veneers (1.0 mm) on the 1.0-mm thick core samples tested in tension exhibited the lowest strength (158 MPa), whereas the thinner veneer porcelain (0.5 mm) on the 1.5-mm thick core samples tested for compressive strength recorded substantially greater strength (415 MPa). The superior strength of the Procera AllCeram crown is apparent, but the fit of a restoration to the supporting structure can influence the strength and clinical prognosis of an all-ceramic crown. Tuntiprawon and Wilson13 found that all-ceramic crowns with smaller gap dimensions at the axial wall and marginal opening demonstrated the best compressive strengths.

American Dental Association (ADA) Specification No. 8 states that the luting cement film thickness for a crown restoration should be no more than 25 μm when using a Type I luting agent, or 40 μm with a Type II luting agent.14 Marginal fit of cemented restorations that range from 25 to 40 μm has been suggested as a clinical goal, but marginal openings in the range of these dimensions clinically are seldom achieved. Numerous studies have examined marginal fit of all porcelain crowns (Table I).15, 16, 17, 18, 19, 20, 21, 22, 23, 25, 26, 27, 28, 29, 30, 31, 32, 33, 34, 35, 36, 39 and 40After examining more than 1000 crowns clinically after a 5-year period, McLean and von Fraunhofer37 and McLean38 concluded that a marginal opening of ≤120 μm was clinically acceptable. The aluminous porcelain jacket crown with the Vita-Vitadur N core porcelain recorded marginal openings that ranged from 0 to 36 μm.39 Dicor ceramic crown reported marginal openings that ranged from 16 to 120 μm,20, 21, 25, 26, 27, 28 and 29 and Cerestore ceramic crown marginal openings ranged from 1.0 to 313 μm,20, 21, 25 and 26 whereas the Renaissance/Ceplatec ceramic System20 and 27 ranged from 20 to 76 μm. Marginal openings reported for the Empress ceramic crown system were 55 μm SD ± 29 μm,36 whereas the In-Ceram ceramic system demonstrated the overall largest range (34 to 123 μm).33 and 40 Collectively, studies of all-ceramic crown systems have reported mean marginal openings of approximately 155 μm with a range of 0 to 313 μm.15, 16, 17, 18, 19, 20, 21, 22, 23, 25, 26, 27, 28, 29, 30, 31, 32, 33, 34, 35, 36, 39 and 40 Based on these reports, it would appear that 100 μm is a suitable measurement value for a clinically acceptable crown.

Table I.

The marginal opening of various all-ceramic crown systems reported in the literature

| System | Marginal opening (MO) | References |
| --- | --- | --- |
| Aluminous porcelain jacket (Vita-Vitadur N core porcelain) | 0.0 - 36.0 μm | 39 |
| Dicor | 16.0 - 120.0 μm | 20, 21, 25, 26, 27, 28 and 29 |
| Cerestore | 1.0 - 313.0 μm | 20, 21, 25 and 26 |
| Renaissance/Ceplatec | 20.0 - 76.0 μm | 20 and 27 |
| Empress | 55.0 (±29) μm | 33 and 41 |
| In-Ceram | 34 - 123 μm | 33 and 41 |

Attanasi41 evaluated color stability of Procera AllCeram coping and veneering porcelain in an in vitro study under conditions of accelerated aging. After 1200 hours of accelerated aging, the color change between the pretest and post test samples, measured as ΔE values with the Chroma Meter II instrument, was 1.41 ± 1.10 for 300 hours and 1.03 ± 0.62 for 1200 hours. None of the Procera samples demonstrated a color change that could be detected clinically.

Measurement data for the marginal fit of the Procera AllCeram crown has not been reported. Therefore the purpose of this investigation was to (1) measure the marginal opening and internal adaptation of Procera AllCeram crowns fabricated for premolar and molar teeth to the working dies, and (2) determine the precision of fit between the crowns and the dies achieved by the Procera CAD/CAM system for the fabrication of all-ceramic restorations.

- MATERIAL AND METHODS

The upper member of a Kilgore typodont (Model 500-B, Kilgore International Incorporated, Coldwater, Mich.) was used to fabricate 5 patient simulation models of a maxillary arch. Five ivorine maxillary right firstpremolars and first molars were prepared with chamfer margins for complete crowns. Tooth preparations were standardized with: (1) a total convergence angle of 10 degrees, (2) chamfer margins 1.3 to 1.5 mm circumferentially, and (3) occlusal reduction of 2.0 mm. All line-angles were rounded. One premolar and 1 molar were placed in their correct positions in the maxillary arch of the typodont. A full arch impression was made of the preparations and adjacent teeth with polyvinyl siloxane impression material (Extrude, Light and Heavy bodied, Kerr Mfg., Romulus, Mich.) in a custom tray with vertical stops and according to manufacturer’s instructions. After the prescribed setting period, the impression was removed from the upper member of the typodont. The prepared ivorine teeth were removed from the typodont and inserted into the impression at the correct location and orientation. This impression with the prepared ivorine teeth was poured in an improved dental stone to create a patient simulation model. Five patient simulation models were created as the experimental population with this method. In turn, a second impression was made of each simulation model to fabricate master casts with removable premolar and molar dies for each patient. Final impressions of the simulation models were poured in an epoxy resin material (Epoxy potting resin, General Polymers Corp., Cincinnati, Ohio). After the resin set, pins (Pindex System, Whaledent Int., New York, N.Y.) were placed in the appropriate locations, and the base of the cast was poured in dental stone. Dies were removed from the cast base, and trimmed to the preparation margins. Five master casts with removable dies were fabricated for the 5 patient simulations with this technique (Fig. 2, Aand B).

|  | Display Full Size version of this image (109K) |
| --- | --- |

Fig. 2. A, Dies of premolar and molar teeth. B, Five sets of dies were fabricated for each of 5 patient simulation models.

Procera CAD/CAM technology was used to fabricate all-ceramic restorations. In this system, the die isdigitized by a scanning instrument and the data for the tooth preparation are then transmitted electronically to the laboratory (Procera Sandvik AB, Stockholm, Sweden). At the laboratory, data are used to fabricate the sintered aluminous oxide coping.10 After receiving the coping from the laboratory, it is returned to its respective die and the anatomic features of the tooth are built-up by the ceramic technician with Procera AllCeram low fusing surface veneering porcelain (Fig. 3, A and B).

|  | Display Full Size version of this image (132K) |
| --- | --- |

Fig. 3. A, Procera densely sintered all-ceramic core for molar tooth. B, Procera AllCeram crown after low fusing surface porcelain is added to core to restore anatomic form of tooth (premolar).

Laser videography (Chesapeake Laser System Model LTG-2105 Lanham, Md.) was used to measure the precision of fit between the Procera AllCeram crowns and the dies. This measurement technique uses a laser to digitize sample structure and to collect 3-dimensional data. A Gallium Arsenide laser is the optic source for the system that operates at a wavelength of 780 nm with the capacity to measure at the micrometer level. The reported accuracy of this laser instrument is ± 1.0 μm.

The measurement protocol uses 3 spheres that function as references in matching the die surface to the internal surface of the crown. To measure the precision of fit of a crown to the die, the crown is positioned on the die with a silicone impression material as the luting agent (Extrude Light-bodied, Kerr Mfg. Co.). A thin lay of silicone is placed within the crown, which is seated on the die and held under a load of 1.36 kg for 10 minutes. When the silicone is set, the crown is carefully removed from the die, leaving the silicone intact with the die (Fig. 4).

|  | Display Full Size version of this image (52K) |
| --- | --- |

Fig. 4. Intact silicone covering die provided 3-dimensional topography of internal aspects of Procera crown. 
Thus, silicone provides: (1) the physical dimension of the 3-dimensional space between the crown and the die, and (2) the topography of the internal surface of the crown.

A 12.0 × 12.0 mm area that completely covered the silicone material and die was digitized by using an x-axis and y-axis measurement matrix of 100 μm between each data point. The surface of the reference spheres was also digitized and their 3-dimensional relationships to the die were recorded. When digitizing was completed, the silicone luting agent was removed from the die and the die was digitized without removal from the digitizer. The database created by this procedure provided measurements of the precision of fit for the Procera AllCeram crowns (Fig. 5, A and B).

|  | Display Full Size version of this image (128K) |
| --- | --- |

Fig. 5. A, Graphic replication of internal aspects of Procera AllCeram crown. B, Graphic illustration of digitized database collected by laser videography.

The laser videography measurement system, through its computer software, is capable of matching the surface of the silicone (internal surface of crown) to the surface data of the die (tooth preparation) by using the 3 spheres as references. Once matched, the system is able to select from the matched crown-die database any number of cross-sectional scan lines in either the facial or the proximal directions. In turn, the scan lines are used to determine the precision of fit of the crown to the die. From these 2-dimensional scans, the marginal discrepancies and internal gaps can be measured at multiple locations between the die and crown similar to what has been previously reported.42 and 43 For each crown, the following 4 measurement locations (Fig. 6)were used to determine the precision of fit between the crowns and the dies for the premolar and the molar teeth:

1. Marginal Open (MO): The marginal opening at the point of closest approximation between the epoxy die and porcelain margin of the crown.

2. Axial Wall (AW): The internal adaptation of the crown walls at the midpoint of the axial wall (2.0 mm occlusal to the margin of the die).

3. Cusp Tip (CT): The internal adaptation at the cusp tips.

4. Occlusal Adaptation (OA): The internal adaptation of the occlusal surface of the crown to the die at the midpoint from the facial and proximal.

|  | Display Full Size version of this image (56K) |
| --- | --- |

Fig. 6. Crown to die diagram showing measurement locations along scan line used in sampling of database to determine marginal opening, internal adaptation, and precision of fit at the crown to die interface.

The fit of the Procera AllCeram crowns to the epoxy dies was evaluated. By using the scan line scheme planned for the investigation, measurements were taken from the database at the MO, AW, CT, and OA measurement locations to evaluate the fit of all crown samples. Figure 6 shows these measurement locations along a single scan line through a section of a crown-die sample. The premolar required 9 scan lines viewed from the facial and proximal surfaces that were separated by 500 μm and by 200 μm, respectively. The molar scan lines viewed in both the facial and the proximal directions were separated by 700 μm. Data recorded at the different scan lines of the facial and proximal scan line directions were averaged by the different teeth (premolar and molar) and measurement locations (MO, AW, CT, and OA).

Mean data were calculated and analyzed statistically with descriptive statistics and repeated measurement analysis44 to assess the between-crown groups effect, within-crown group measurement locations effect, and the interactions between-crown groups and within-crown groups. The Greenhouse-Geisser epsilon (G-G epsilon) adjustment was performed as part of the repeated measure analysis with the SuperANOVA program.45 Data analyses for the investigation were performed at .05 level of significance with the statistical program SuperANOVA (Abacus Concepts, Berkeley, Calif.).

The average gap dimension at each measurement location within each crown group was determined by calculating the means and standard deviations (SD). Internal adaptation was determined by calculating the mean and SDs of the average gap dimensions for the internal measurement locations by crown group. Precision of fit was determined by using the average gap dimension of all the measurement locations to calculate the mean and SD by crown group. Repeated measurement analyses were performed for the between-crown factor (premolar/molar), the within-crown factor measurement locations (MO, AW, CT, and OA), and the interactions of these 2 factors.

The G-G epsilon is an adjusted probability for all the within-subjects hypothesis tests. The epsilons were designed to help eliminate the effects of correlation within the repeated measurement on the validity of the tests. The procedure used to measure the gap at the measurement locations of the crown-die interface were considered repeated measurements within-crown groups on the same experimental unit and were correlated.

- RESULTS

Figure 7 shows the means and SDs of the average gap dimensions of the crown groups by measurement locations.

|  | Display Full Size version of this image (38K) |
| --- | --- |

Fig. 7. Mean gap dimension and standard deviation of crown groups at marginal opening (MO), axial wall (AW), cusp tip (CT), and occlusal adaptation (OA) measurement locations for premolar and molar crowns. 
The MO gap dimension mean and SD for the premolar and molar crowns were 56.0 μm SD ± 21 μm and 63.0 μm SD ± 13 μm, respectively. The internal adaptation gap dimensions means and SDs for the premolar crowns were 69.0 μm SD ± 17 μm (AW), 48.0 μm SD ± 12 μm (CT), and 36.0 μm SD ± 7 μm (OA). Although the internal adaptation gap dimensions means and SDs for the molar crowns were 49.0 μm SD ± 3 μm (AW), 67.0 μm SD ± 21 μm (CT), and 74.0 μm SD ± 29 μm (OA), the precision of fit gap dimension means and SDs for the premolar and molarcrowns were 52.0 μm SD ± 19 μm and 63 μm SD ± 20 μm, respectively (Fig. 8).

|  | Display Full Size version of this image (51K) |
| --- | --- |

Fig. 8. Mean gap dimension and standard deviation of crown to die precision of fit for premolar and molar crowns. 
The precision of fit gap dimensions for the crown groups were equal to or less than the crown groups mean MO gap dimension.

Table II shows the analysis of variance (ANOVA) of the repeated measurement analyses.
Table II.

Repeated measures analysis by between-crown factor (premolar/molar), within-measurement location factor (MO, AW, CT, and OA), and the interaction of the 2 factors

| Source | df | Sum of squares | Mean squares | F value | G-G value |
| --- | --- | --- | --- | --- | --- |
| Crown | 1 | 0.0012 | 0.0012 | 1.8934 |  |
| Subject (group) | 8 | 0.0052 | 0.0007 |  |  |
| Measurement location | 3 | 0.0001 | 4.6733E-5 | 0.2647 | 0.8098 |
| Measurement location * crown | 3 | 0.0044 | 0.0015 | 8.3687 | 0.0015* |
| Measurement location * subject (crown) | 24 | 0.0042 | 0.0002 |  |  |
| G-G Epsilon |  |  |  |  |  |
| åMeasurement location |  | 0.8082 |  |  |  |
| *Significant at .05 level. |  |  |  |  |  |

The reported P values were modified by the G-G epsilon. The epsilon factors (å) for the degree of freedom (df) adjustment are provided in the lower left. The probabilities are not corrected for values of epsilon greater than 1.

There were no significant differences found between the crown groups gap dimensions or within crown groups measurement locations gap dimensions (G-G=0.8098). The interactions of the 2 factors were significant (G-G=0.0015).

To identify the source of the significant interactions, data were further analyzed. Because the repeated measure analysis contained both within and between factors, the pairwise comparisons will not be allowed. Therefore 1-way ANOVA46 on the crown was performed. Table III presents the 1-way ANOVAs of the crown groups by MO, AW, CT, and OA measurement locations.Mean gap dimensions between the crown groups were significantly different at AW and OA measurement locations. MO and CT measurement locations mean gap dimensions did not demonstrate any significant differences between crown groups. To conduct comparison analysis of the within crown measurement locations gap dimensions, the data were separated by crown group and analyzed with 1-within factor repeated measures ANOVA by crown factor (premolar/molar).

Table III.

One-way ANOVAs of between-crown factor by measurement locations (MO, AW, CT, and OA)

| Source |  | df | Sum of squares | Mean squares | F value | P value |
| --- | --- | --- | --- | --- | --- | --- |
| MO | Crown | 1 | 0.0001 | 0.0001 | 0.3812 | .5541 |
|  | Residual | 8 | 0.0024 | 0.0003 |  |  |
| AW | Crown | 1 | 0.0010 | 0.0010 | 6.6000 | .0332* |
|  | Residual | 8 | 0.0012 | 0.0001 |  |  |
| CT | Crown | 1 | 0.0009 | 0.0009 | 3.0669 | .1180 |
|  | Residual | 8 | 0.0024 | 0.0003 |  |  |
| OA | Crown | 1 | 0.0036 | 0.0036 | 8.5186 | .0193* |
|  | Residual | 8 | 0.0034 | 0.0004 |  |  |

Significant difference at .05 level.

The ANOVA of the repeated measure analysis for the premolar crowns within-measurement location factor (Table IV) revealed significant differences in the measurement locations gap dimensions (G-G= 0.0293), whereas the molar crowns displayed no significant differences in the measurement locations gap dimensions (Table V).However, measurement location gap dimensions were approaching a marginal level of significant differences (G-G=0.1423). The interactions between measurement location and crown were not significant for either crown group.

Table IV.

One-within factor (premolar) repeated measures ANOVA on the measurement location factor (MO, AW, CT, and OA).

| Source | df | Sum of squares | Mean squares | F value | G-G value |
| --- | --- | --- | --- | --- | --- |
| Premolar | 4 | 0.0016 | 0.0004 |  |  |
| Measurement location | 3 | 0.0030 | 0.0010 | 5.7164 | 0.0293* |
| Measurement location * premolar | 12 | 0.0021 | 0.0015 |  |  |
| G-G Epsilon |  |  |  |  |  |
| åMeasurement location |  | 0.6591 |  |  |  |

Significant at .05 level.

Reported P value was modified by the G-G epsilon. The epsilon factors (å) for the degree of freedom (df) adjustment are provided in the lower left. The probabilities are not corrected for values of epsilon greater than 1.

Table V.

One-within factor (molar) repeated measures ANOVA on measurement location factor (MO, AW, CT, and OA)

| Source | df | Sum of squares | Mean squares | F value | G-G value |
| --- | --- | --- | --- | --- | --- |
| Molar | 4 | 0.0036 | 0.0009 |  |  |
| Measurement location | 3 | 0.0016 | 0.0005 | 2.9492 | 0.1423 |
| Measurement location * molar | 12 | 0.0021 | 0.0002 |  |  |
| G-G Epsilon |  |  |  |  |  |
| åMeasurement location |  | 0.4406 |  |  |  |

Reported P value was modified by G-G epsilon. Epsilon factors (å) for the degree of freedom (df) adjustment are provided in the lower left. Probabilities are not corrected for values of epsilon greater than 1.

To compare the within crown measurement locations gap dimensions, the contrast test47 was performed at the .05 level of significance. Three measurement locations had significant differences in gap dimensions for the premolar crowns (Table VI).The molar crowns had 1 measurement location with a gapdimension of significant difference (G-G=0.0423). The AW and OA gap dimensions were approaching marginal significance (G-G = 0.0960) in the molar crowns.

Table VI.

Contrast comparison test of crown factor (premolar/molar) for the effect by measurement location factor (MO, AW, CT, and OA)

| Source |  | df | Sum of squares | Mean squares | F value | G-G value |
| --- | --- | --- | --- | --- | --- | --- |
| Premolar | MO vs AW | 1 | 0.0004 | 0.0004 | 2.5721 | 0.1466 |
|  | vs CT | 1 | 0.0002 | 0.0002 | 0.9632 | 0.3112 |
|  | vs OA | 1 | 0.0010 | 0.0010 | 5.8450 | 0.0512* |
|  | AW vs CT | 1 | 0.0012 | 0.0012 | 6.6833 | 0.0411* |
|  | vs OA | 1 | 0.0028 | 0.0028 | 16.6833 | 0.0065* |
|  | CT vs OA | 1 | 0.0004 | 0.0004 | 2.0627 | 0.1805 |
| Molar | MO vs AW | 1 | 0.0004 | 0.0004 | 2.4384 | 0.1563 |
|  | vs CT | 1 | 4.4100E-5 | 4.4100E-5 | 0.2469 | 0.4481 |
|  | vs OA | 1 | 0.0003 | 0.0003 | 1.7555 | 0.1971 |
|  | AW vs CT | 1 | 0.0008 | 0.0008 | 4.2370 | 0.0960 |
|  | vs OA | 1 | 0.0015 | 0.0015 | 8.3318 | 0.0432* |
|  | CT vs OA | 1 | 0.0001 | 0.0001 | 0.6857 | 0.3210 |
| *Significant at .05 level. |  |  |  |  |  |  |

The reported P values were modified by the G-G epsilon. The cell weights used in the contrast test were from plus (+)1.0000 to minus (–)1.0000.

- DISCUSSION

An acceptable MO for full crowns, as reported by Hung et al.,26 is 50 to 75 μm, whereas Weaver et al.29 suggested 70 μm SD ± 10 μm. The MO mean gap dimensions for the premolar and molar crowns groups in this investigation were 56.0 μm SD ± 21 μm and 63.0 μm SD ± 13 μm, respectively. Many available all-ceramic crown systems demonstrate much larger MO gap dimensions (Table I). The MOs of the all-ceramic crown system tested in this investigation are comparable and, in some cases, better than the gap dimensions reported by other systems.17, 20, 21, 25, 26, 27, 28, 29, 30, 31, 32, 33, 34, 35, 36, 37, 38, 40 and 41 Mean MOs between the crown groups in this investigation were not significantly different.

The absolute marginal opening (AMO), as described by Holmes et al.,42 can be demonstrated when a line connects the “A and E” reference points in Figure 6. The AMO is reported in most studies of crown fit using MO as a generic term. The AMO was not used in this study, because it was the consensus of the investigators that the AMO would vary significantly, depending on the overextension or underextension of the crown sample. The MO, the opening demonstrated when a line connects the “B and E” reference points, would be unaffected by overextension or underextension of the crowns (Fig. 6). Clinically, it is important to know how well a crown fit the tooth at the margin, but equally important is the dimension of the luting space.13 The cement space or internal adaptation is considered to be a uniform space that facilitates seating without compromising retention and resistance forms. This is of paramount importance because all-ceramic crowns are more fragile than the cast alloy alternatives. The precision of fit of the crown can influence the clinical prognosis.Tuntiprawon and Wilson13 reported that all-ceramic crowns displayed greater compressive strength when the mean AW was at a gap dimension of 73.0 μm. Their study also showed that when the mean AW was increased to 122.0 μm, a lower failure strength occurred without any significant improvement in seating. Therefore an all-ceramic crown that can be predictably fabricated with a precision of fit and does not compromise the tooth has the best prognosis for clinical success. Our findings show that the crowns studied will consistently seat with an MO gap dimension of lesser magnitude than others have shown to be associated with optimal all-ceramic crown strength.13 A gap dimension increase of 49 μm above the ideal gap (73.0 μm) for the AW dimension only decreased the mean MO gap dimension approximately 3.0 μm.13 Therefore an acceptable AW gap dimension should be within 1 SD of the ideal AW dimension. The mean AW gap dimension for the premolar crowns (69.0 μm SD ± 17 μm) was not significantly different from the literature report, but the AW for the molar was 23.0 μm less (49.0 μm SD ± 3 μm). The data do not indicate that there were incidences of axial wall contact between the dies and the crowns because contact between the axial walls would cause impingement and affect seating. Nevertheless, the mean gap dimensions at all measurement locations were less than the mean gap dimensions shown to be associated with higher failure strength of the porcelain jacket crown.13 Thus, when precision of fit is combined with strength11, 12 and 13 and color stability,42 the clinical prognosis is enhanced.

The 1-way ANOVAs (Table III) revealed that mean gap dimensions between crown groups were significantly different for AW and OA measurement locations. Figure 7 illustrates the inverse relationship the measurement locations mean gap dimensions have between crown groups. The premolar clearly has a greater gap dimension than the molar at the AW measurement location. Yet, no significant differences exist in gap dimensions at the CT measurement locations between crown groups, but the molar gap dimension was greater than the premolar. Gap dimensions at the OA measurement locations between crown groups were significantly different, and the molar gap dimension was greater than the premolar. Premolar gap dimensions decreased from AW>CT>OA, whereas the molar gap dimensions increased from AW<CT<OA between crown groups.

A luting space of 49 μm at the AW for the molar was much less than the 69 μm for the premolar. The larger AW gap dimension in the premolar group combined with a smaller MO gap dimension suggest that further seating could possibly occur. The larger gap dimensions at the CT (67 μm) and OA (74 μm) locations for the molar when compared with the premolars CT (48 μm) and OA (36 μm) also suggest incomplete seating. However, mean CT gap dimension between crown groups does not support incomplete seating. Statistically, a 19 μm difference between the crown groups CT gap dimensions was not significant, whereas a 38 μm difference at the OA was significant. Theoretically, a decrease of 19 μm at the CT in the molar group would still leave a gap dimension of 19 μm at the OA. An improvement of this magnitude in the gap dimension for the MO and AW measurement locations would not be feasible. The analysis shows a between crown group difference in mean MOs of 7.0 μm. If a gap reduction of 19 μm occurred at the CT as previously stated, the difference in gap dimensions between the crown groups MOs should not change the statistical inference. This might lead one to question whether the differences at AW, CT, and OA were actual differences in the fabrication of the crowns, or whether the differences were caused by the measurement method. At best, the overall difference in the crown groups MOs would be 12.0 μm with the hypothetical 19.0 μm improvement. Gap dimensions at the MO and CT measurement locations were comparable between crown groups. If there were errors in the method used to measure the fit, then the database would be composed of inflated values and the fit of the crowns would be much better than our reported findings. The most important question to address is whether the differences are significant clinically. Obviously, the AW and OA can influence the hydraulic pressure created during cementation procedures regardless of the viscosity of the cement mixture. If the luting agent and method used in this study was the cause of the significant differences recorded in the gap dimensions, the fit of the crowns is still an improvement in technology. The MO mean gap dimensions for crown groups were comparable to previous reports.18, 21, 22, 26, 27, 28, 29, 30, 31, 32, 33, 34, 35, 36, 37, 38, 39, 41 and 42 The internal adaptation and precision of fit for the crowns were comparable to the MO gap dimensions.

The 3 most common instruments used to measure the fit of the crowns to the supporting structure are stereomicroscopy, traveling microscopy, and videomicroscopy. The method used with most of these instruments provide limited databases generated from widely separated data points. The measurements recorded with these instruments are taken at a single location per axial surface at the margin of the crown. The mean dimension calculated for a specific measurement location, for example, the MO, have often demonstrated excessive standard deviations. Subsequently, questions have been raised regarding the repeatability of these methods and the adequacy of the database used in the documentation of the crowns fit. Chan et al.24 and Karlsson25 reported differences in the MOs between a crown and die in the range of 54 to 130 μm within a distance of 300 μm along margins of the restorations. The laser method used in our study generated an average of 8,000 and 12,000 data points for the premolar and molar dies, respectively. From this database, 9 scan lines in the facial and proximal directions were selected for each matched crown to die sample. From the 18 scan lines, 36 measurements were made for each MO, AW, and CT measurement location, whereas 18 measurements were recorded for the OA measurement location for a total of 126 measurements per crown and die sample. Thus, a total of 252 points per sample were used to analyze the fit of each crown. The repeated measures ANOVA showed no significant differences in the average gap dimension for the 4 measurement locations between the crown groups, and within crown group measurement locations gap dimensions had no significant effect. The significant interaction (G-G=0.0015) of the measurement location gap dimension by crown is of interest because it indicates that the differences in measurement location gap dimension is dependent on the crown group and the location.

The separate within-factor repeated measures ANOVAs indicated that the premolar measurement location gap dimensions were significantly different (G-G=0.0293, Table IV), whereas the molars were not significantly different (G-G=0.1423, Table V). Although the differences in the measurement locations gap dimensions for the molars were not significant at the level of .05, maybe a significant difference would occur with a larger sample size or with a cement of a different viscosity (Table V).

To isolate the significant differences in gap dimensions within crown groups, separate contrast comparison tests were performed by the measurement locations. The within-crown group differences in gap dimensions were clearly demonstrated (Table VI). The smaller crowns had significant differences in measurement locations gap dimensions when comparing the MO to OA, the AW to CT, and the AW to OA, whereas the larger crowns demonstrated significant differences only in the AW to OA measurement locations gap dimensions. These findings further support the position that the differences in the gap dimensions are dependent on the crown and the measurement location.

Another possible explanation for the difference within crown groups could be related to the refractory die fabrication process. The difference in the calculation and the ability of the Procera CAD/CAM to perform at the level of accuracy required to fabricate the refractory die in the suggested 15% to 20% enlargement range9 and 10 may be the cause of these differences. This is supported by the findings that there were more significant differences in the gap dimensions of the measurement locations of the crowns with the smallest surface area. This requires further study, beyond the scope of this investigation, to determine the cause of these significant differences.

Mean gap dimensions for the MO measurement location in Figure 1 is of particular interest to clinicians because it represents the gap between the external surface of the tooth and the crown margins. It is paramount that the MO be kept to a minimum to establish clinical acceptability of the fit for any crown system. Ideally this space should be in the 25 to 40 μm range to accommodate the different types of luting agents, but most often the MOs are greater. In our study, the luting space was recorded by positioning the crowns on the respective dies with a light-bodied silicone to simulate the luting agent. This approach has been used by other investigators in measuring the gap dimension.27, 37 and 38 McLean and von Fraunhofer38 demonstrated that cementation simulation with polyether rubber will require a space of 22 μm, whereas 20 μm is needed for zinc-phosphate cement. Therefore it is theorized that the luting space measured and represented by the silicone impression did not prevent the accurate seating of the crown as a result of hydraulic pressure. Furthermore, data analysis confirmed that no impingements occurred at the measurement locations used. The available space is more than enough to accommodate the space required for the commonly used Type I and II luting agents of approximately 25 to 40 μm.

- CONCLUSIONS

This study used an in vitro model and ivorine teeth and a light body impression material functioning in a dual role to simulate cement and to capture the accurate topography of the internal aspect of the crowns. The following conclusions were drawn.

1. Quantitatively generated database of the precision of fit at the crown-die interface using laser videography combined with the provisional cementation technique appears to be a valid method to quantify the marginal gap, cement space, or to determine the precision of fit at the crown-die interface.

2. The Procera AllCeram crown for the posterior dentition demonstrated a marginal gap at the crown-die interface that was less than 70 μm.

3. Precision of fit at the crown-die interface for premolar and molar crowns was 56 ± 21 μm and 63 ± 13 μm, respectively.

4. The difference in the fit between crown groups depend on the region of the crown being evaluated and the type of crown.

5. The Procera CAD/CAM system consistently produced all-ceramic crowns for premolar and molar teeth with a clinically acceptable fit at the margin (54 to 64 μm) and internal aspects (49 to 63 μm), which is conducive to clinical success.

- Acknowledgements

We would like to acknowledge the assistance of Beth E. Lang in the operation of the laser digitizer and collection of data. We also thank Mr. Rui-Feng Wang for his assistance in the preparation of the measurement data and the performance of the several statistical tests used in this investigation

- References

1 Cerestore crown technical manual. Fabrication process. Lakewood (CO): Coors Biomedical Co.

2 Dicor laboratory technical manual, Dentsply International Inc, York (1987).

3 KH. Koerber, Ceraplatin crown structures with reproducible minimal thickness strength, Huethig, Heidelberg (1985).

4 JG Wall and DL. Cipra, Alternative crown systems. Is the metal-ceramic crown always the restoration of choice?, Dent Clin North Am 36 (1992), pp. 765–782. View Record in Scopus | Cited By in Scopus (12)

5 JW McLean and TH. Hughes, The reinforcement of dental porcelain with ceramic oxides, Br Dent J 119 (1965), pp. 251–267.

6 S. Timoshenio, Theory of plates and shells, McGraw-Hill, New York (1959) 79.

7 SF Hulbert, JC Brokros, LL Hench, J Wilson and G. Heimke, Ceramics in clinical applications, past, present and future. In: P Vincenzinc, Editor, High tech ceramics., Elsevier Science, Amsterdam (1987), pp. 3–27.

8 P. Boutin, Total arthroplasty of the hip by fitted aluminum prosthesis. Experimental study and 1st clinical application, Rev Chir Orthop Appar Mot 58 (1972), pp. 229–246 [in French]. View Record in Scopus | Cited By in Scopus (99)

9 M Andersson and A. OdÉn, A new all-ceramic crown: a dense-sintered, high-purity alumina coping with porcelain, Acta Odontol Scand 51 (1993), pp. 59–64. View Record in Scopus | Cited By in Scopus (90)

10 M Andersson, B Bergman, C Bessing, G Ericson, P Lundquist and H. Nilsson, Clinical results with titanium crowns fabricated with machine duplication and spark erosion, Acta Odontol Scand 47 (1989), pp. 279–286. View Record in Scopus | Cited By in Scopus (57)

11 WC Wagner, TM Chu and ME. Razzoog, Indentation fracture toughness of three new ceramic core materials, J Dent Res 74 (1995), p. 471 (Abstract 568). Abstract | PDF (1118 K) | View Record in Scopus | Cited By in Scopus (11)

12 WC Wagner and TM. Chu, Biaxial flexural strength and indentation fracture toughness of three dental core ceramics, J Prosthet Dent 76 (1996), pp. 140–144. SummaryPlus | Full Text + Links | PDF (1212 K) | View Record in Scopus | Cited By in Scopus (69)

13 M Tuntiprawon and PR. Wilson, The effect of cement thickness on the fracture strength of all-ceramic crowns, Aust Dent J 40 (1995), pp. 17–21. View Record in Scopus | Cited By in Scopus (17)

14 American Dental Association, ANSI/ADA Specification No. 8 for zinc phosphate cement. In: Guide to dental materials and devices (5th ed), American Dental Association, Chicago (1970–71).

15 PJ Adair and DG. Grossmann, The castable ceramic crown, Int J Periodont Restor Dent 4 (1984), pp. 33–45.

16 C Chan, G Haraszthy, J Geis-Gerstorfer and H. Weber, The marginal fit of Cerestore full-ceramic crowns—a preliminary report, Quintessence Int 6 (1985), pp. 399–402.

17 DG. Grossman, Cast glass ceramics, Dent Clin North Am 29 (1985), pp. 725–739. View Record in Scopus | Cited By in Scopus (21)

18 T Sato, A Wohlend and P. Schärer, “Shrink-free” ceramic crown system: factors influencing the core marginal fit, Quintessence Dent Technol 10 (1986), pp. 81–86.

19 CE Brukl and GK. Philp, The fit of molded all-ceramic, twin foil, and conventional ceramic crowns, J Prosthet Dent 58 (1987), pp. 408–413. Abstract | View Record in Scopus | Cited By in Scopus (3)

20 DR. Davis, Comparison of fit of two types of all-ceramic crowns, J Prosthet Dent 59 (1988), pp. 12–26.

21 P Schärer and T. Sato, A comparison of the marginal fit of three cast ceramic crown systems, J Prosthet Dent 59 (1988), pp. 534–542.

22 A Wohlwend, JR Strub and P. Schärer, Metal ceramic and all-porcelain restorations: current considerations, Int J Prosthodont 2 (1989), pp. 13–26. View Record in Scopus | Cited By in Scopus (17)

23 C Chan, G Haraszthy, J Geis-Gerstorfer, H Weber and H. Huettemann, Scanning electron microscopic studies of the marginal fit of three esthetic crowns, Quintessence Int 20 (1989), pp. 189–193. View Record in Scopus | Cited By in Scopus (9)

24 S. Karlsson, The fit of Procera titanium crowns. An in vitro and clinical study, Acta Odontol Scand 51 (1993), pp. 129–134. View Record in Scopus | Cited By in Scopus (31)

25 MF Abbate, AHL Tjan and WM. Fox, Comparison of the marginal fit of various ceramic crown systems, J Prosthet Dent 61 (1989), pp. 527–531. Abstract | View Record in Scopus | Cited By in Scopus (29)

26 SH Hung, KS Hung, JD Eick and RP. Chappell, Marginal fit of porcelain-fused-to-metal and two types of ceramic crowns, J Prosthet Dent 63 (1990), pp. 26–31. Abstract | View Record in Scopus | Cited By in Scopus (12)

27 F Vahidi, ET Egloff and FV. Panno, Evaluation of marginal adaptation of all-ceramic crowns and metal ceramic crowns, J Prosthet Dent 66 (1991), pp. 426–431. Abstract | View Record in Scopus | Cited By in Scopus (3)

28 KM Harrison, EJ Billy and GB. Pelleu, Effects of an internal escape channel on a castable ceramic crown, J Prosthet Dent 65 (1991), pp. 622–627. Abstract | View Record in Scopus | Cited By in Scopus (0)

29 JD Weaver, GH Johnson and DJ. Bales, Marginal adaptation of castable ceramic crowns, J Prosthet Dent 66 (1991), pp. 747–753. Abstract | View Record in Scopus | Cited By in Scopus (13)

30 T Hummert, N Barghi and T. Berry, Post cementation marginal fit of a new ceramic foil crown system, J Prosthet Dent 68 (1992), pp. 766–770. Abstract | View Record in Scopus | Cited By in Scopus (2)

31 A Lomanto and S. Weiner, A comparative study of ceramic crown margins constructed using different techniques, J Prosthet Dent 67 (1992), pp. 773–777. Abstract | View Record in Scopus | Cited By in Scopus (0)

32 RJ Hoard, PC Chiang, ER Hewlett and AA. Caputo, Marginal discrepancy as related to margin design in porcelain-fused-to-Dicor restorations, Oral Health 83 (1993), pp. 15–16.

33 NJ Grey, V Piddock and MA. Wilson, In-vitro comparison of conventional crowns and a new all-ceramic system, J Dent 21 (1993), pp. 47–51. Abstract | View Record in Scopus | Cited By in Scopus (20)

34 AJ Schonenberger, A Di Felice and M. Cossu, Improving the precision of esthetic ceramic margins: guidelines for success, J Esthet Dent 6 (1994), pp. 143–150. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (0)

35 P Pera, S Gilodi, F Bassi and S. Carossa, In vitro marginal adaptation of alumina porcelain ceramic crowns, J Prosthet Dent 72 (1994), pp. 585–590. SummaryPlus | Full Text + Links | PDF (2114 K) | View Record in Scopus | Cited By in Scopus (18)

36 MI Fanuscu and JA. Sorensen, Reduced microleakage of adhesive ceramic crowns by modified cementation methods. IADR, J Dent Res 72 (1993), p. 186 (Abstract 663) (Special Issue).

37 JW McLean and JA. von Fraunhofer, The estimation of cement film thickness by an in vivo technique, Br Dent J 131 (1971), pp. 107–111. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (68)

38 JW. McLean, Polycarboxylate cements. Five years' experience in general practice, Br Dent J 132 (1972), pp. 9–15. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (1)

39 AJ Dickinson, BK Moore, RK Harris and RW. Dykema, A comparative study of the strength of aluminous porcelain and all-ceramic crowns, J Prosthet Dent 61 (1989), pp. 297–304. Abstract | View Record in Scopus | Cited By in Scopus (13)

40 S Rinke, G Margraf, L Jahn and A. Huls, The quality appraisal of copy-milling complete-ceramic crown structures (Celay/In-Ceram), Schweiz Monatsschr Zahnmed 104 (1994), pp. 1495–1499 [in German]. View Record in Scopus | Cited By in Scopus (5)

41 RC. Attanasi Jr, Color stability of Procera all-ceramic porcelain. [Masters thesis.]. In: H Horace, Editor, Rackham School of Graduate Studies., University of Michigan, Ann Arbor (1995).

42 S Valderrama, N Van Roekel, M Andersson, CJ Goodacre and CA. Munoz, A comparison of the marginal and internal adaptation of titanium and gold-platinum-palladium metal ceramic crowns, Int J Prosthodont 8 (1995), pp. 29–37. View Record in Scopus | Cited By in Scopus (19)

43 JR Holmes, SC Bayne, GA Holland and WD. Sulik, Considerations in measurement of marginal fit, J Prosthet Dent 62 (1989), pp. 405–408. Abstract | View Record in Scopus | Cited By in Scopus (42)

44 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 1049-57.

45 BS. Everitt, The Cambridge dictionary of statistics in the medical sciences, Cambridge, Melbourne (1995) 111.

46 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 529-32.

47 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 576-77.

Precisão de Adaptação: Coroa de Procera AllCeram

Kenneth B. maio DDS, MSa, Melinda M. Russell DDS, MSa, Michael E. Razzoog DDS, MS, MPHb e Brien R. Lang DDS, MSc 
Faculdade de Odontologia da Universidade de Michigan Ann Arbor, Mich 

Disponível em linha 4 de Novembro de 2005.

Abstract 
Declaração do problema:

Cor, estabilidade e precisão do término são requisitos para todas as restaurações de cerâmica. O sistema de coroa Procera AllCeram, é densamente sinterizado por alumina e apresenta uma porcelana de superfície com baixa fusão, e parece satisfazer a maioria destes requisitos. No entanto, a avaliação da adaptação marginal não foi relatada.

Finalidade:

Este estudo mede a precisão de adaptação no término das coroas Procera AllCeram fabricadas com a tecnologia Procera CAD / CAM em pré-molares e molares.

Material e métodos:

Cinco primeiros pré-molares e primeiros molares superiores foram preparados para coroa total. Os preparos foram padronizados com uma convergência de um ângulo de 10 graus, término em chanfro, margens de 1,3 a 1,5 milímetros na vestibular, lingual, mesial e distal e, redução oclusal de 2,0 milímetros. Foram avaliados os términos dos preparos das coroas AllCeram através de um procedimento padronizado com uma impressão com um material a base de silicone que teve uma dupla função: (1) preencher o espaço do agente de cimentação, e (2) para replicar os aspectos internos da coroa. Laser videografia foi utilizado para medir a diferença entre a dimensão e as desadaptações marginais das coroas, paredes axiais, as pontas de cúspides, bem como a adaptação oclusal. As dimensões da desadaptação e desvios padrão (EE), foram calculados em relação à desadaptação marginal, adaptação interna, e precisão de ajuste.

Resultados:

As dimensões das desadaptações marginais e desvios padrão nas coroas de pré-molares e molares foram 56,0 µ m SD ± 21 e 63,0 µ m SD ± 13 µ m, respectivamente. A diferença média das dimensões da adaptação interna e EE foram 69,0 µ m SD ± 17 µ m nas paredes axiais, com 48,0 µ m SD ± 12 µ m para ponta de cúspide, e 36,0 µ m SD ± 7 µ m na adaptação oclusal dascoroas de pré-molares; e 49,0 µ m ± 3 µ m SD nas paredes axiais, com 67,0 µ m SD ± 21 µ m nas pontas de cúspides, SD e 74,0 µ m ± 29 µ m na adaptação oclusal de coroas de molares. Precisão de ajuste e EE nas coroas de pré-molares e molares foram 52,0 µ m SD ± 19 µ m e 63,0 µ m SD ± 20 µ m, respectivamente.

Conclusão:

As dimensões da desadaptações marginais, adaptações internas, e a precisão de ajustes das coroas foram inferiores a 70 µ m. Estes resultados mostram que as coroas estudadas podem ser confeccionadas com confiança sabendo que as desadaptações serão sempre inferiores a 70 µ m.

(J Prosthet Dent 1998; 80:394-404.) 


MATERIAL E MÉTODOS 
RESULTADOS 
DISCUSSÃO 
CONCLUSÕES 
Agradecimentos 
Referências 

Os resultados deste estudo in vitro deverá auxiliar o dentista na selecção de um conjunto de cerâmica coroa sistema que tem de marginal fidelidade que é necessário para uma clínica favorável prognóstico. Se as coroas utilizado no presente inquérito são uma representação exacta das copas caber possível no ambiente dinâmico da cavidade oral, então essas coroas podem ser prescritos para a estética e da zona póstero dentição com confiança aceitável que uma clínica caber será alcançada consistentemente . 
A procura de estética dentária restaurações em todos os quadrantes tem feito a todos - cerâmicos restauração um popular restauração. All - cerâmicos restaurações clínicos devem satisfazer exigências de força, a precisão de ajuste, cor e estabilidade para que estes sejam bem sucedidos restorations.1, 2, 3 e 4 Para alcançar a adequada força de um conjunto coroa - cerâmica, de alta pureza de alumina foi Sugerida por McLean e Hughes5 como um material de reforçar a feldspathic porcelana matriz. O reforço de alumina - dentários porcelana levou ao desenvolvimento de vários novos reparadora materials.6, 7 e 8 The Procera AllCeram coroa sistema (Nobel Biocare, Göteborg, Suécia), composto de um densamente sinterizada alta pureza aluminous óxido de lidar com um combinado Baixa fusão superfície porcelana, foi introduzido pela Nobel Biocare AB e Sandvik Hard Materials AB.9 e 10 Este sistema requer 3 etapas de fabricação: (1) criação de um espaço alargado morrem de um grupo de trabalho morrer; (2) sinterização de alta pureza aluminous Óxido de pó em 1550 ° C sobre a morrer para formar um Procera lidar, e (3) preencher o formulário coroa por meio da adição de uma baixa fusão superfície porcelana para a convivência no desenvolvimento da anatomia de um dente (Fig. 1). 
  
Display Full Size versão desta imagem (55 KB) 

Fig. 1. Procera AllCeram coroa restauração. 
Wagner e Chu11 relatadas diferenças substanciais em flexurais força para enfrentar o Procera materiais (687 MPa), quando comparados com o In - Ceram (352 MPa) e Empress (134 MPa) - todos os sistemas cerâmicos reparadora. A força do Procera lidar materiais, combinado com a baixa fusão superfície porcelana, foi testado em tensão e compressão com um biaxial flexurais teste (semelhante a ASTM F394 - 78 e ISO / IDS 6872 testes) .12 A thickest porcelana folheados (1. 0 mm) no de 1,0 mm de espessura núcleo amostras testadas em tensão exibiu a menor resistência (158 MPa), enquanto que o mais fino verniz porcelana (0,5 mm) no a 1.5 - mm de espessura núcleo amostras testadas para compressão força Gravado substancialmente maior pujança (415 MPa). A superioridade força do Procera AllCeram coroa é aparente, mas o ajuste de uma restauração à estrutura de apoio pode influenciar a força eo prognóstico clínico de um conjunto de cerâmica coroa. Tuntiprawon e Wilson13 constatou que todas as coroas de cerâmica com menores dimensões lacuna no axiais parede e marginais abertura demonstrada a melhor compressão fortes. 
American Dental Association (ADA) Specification No. 8 afirma que o luting cimento filme espessura de uma coroa restauração não deve ser superior a 25 µ m quando usando um Tipo I luting agente, ou 40 µ m com um Tipo II luting agent.14 Marginal de caber Cimentado restaurações que variam de 25 a 40 µ m, foi sugerida como uma clínica gol, mas marginais aberturas no intervalo destas dimensões clinicamente raramente são alcançados. Numerosos estudos têm analisado marginais ajuste de todos porcelana coroas (Quadro I) .15, 16, 17, 18, 19, 20, 21, 22, 23, 25, 26, 27, 28, 29, 30, 31, 32, 33 , 34, 35, 36, 39 e 40After examinando mais de 1000 coroas clinicamente depois de 5 anos, McLean e von Fraunhofer37 e McLean38 concluiu que a abertura das marginais ≤ 120 µ m foi clinicamente aceitáveis. O aluminous porcelana casaco coroa com a Vita - Vitadur N núcleo porcelana gravado marginais aberturas que variaram de 0 a 36 µ m.39 Dicor cerâmica coroa relatados marginais aberturas que variaram de 16 a 120 µ m, 20, 21, 25, 26, 27, 28 E 29 e Cerestore cerâmica coroa marginais aberturas variaram de 1,0 a 313 µ m, 20, 21, 25 e 26 que o Renascimento / Ceplatec cerâmica System20 e 27 variou de 20 a 76 µ m Marginal aberturas reportado para a cerâmica Empress coroa sistema foram 55 µ m SD ± 29 µ m, 36 que o In - Ceram cerâmica demonstrou o sistema global de maior alcance (34 a 123 µ m) .33 e 40 Colectivamente, estudos de todos os sistemas têm - cerâmicos coroa Relatados significa marginais aberturas de aproximadamente 155 µ m com um intervalo de 0 a 313 µ m.15, 16, 17, 18, 19, 20, 21, 22, 23, 25, 26, 27, 28, 29, 30, 31, 32 , 33, 34, 35, 36, 39 e 40 Com base nestes relatórios, afigura - se que 100 µ m é uma medida adequada para um valor clinicamente aceitáveis coroa. 
Quadro I. 
A abertura de vários marginais todas - cerâmicos coroa sistemas relatada na literatura 
Sistema Marginal abertura (MO) Referências 
Aluminous porcelana casaco (Vita - Vitadur N núcleo porcelana) 0,0 - 36,0 µ m 39 

Dicor 16,0 - 120,0 µ m 20, 21, 25, 26, 27, 28 e 29 

Cerestore 1,0 - 313,0 µ m 20, 21, 25 e 26 

Renaissance / Ceplatec 20,0 - 76,0 µ m 20 e 27 

Empress 55,0 (± 29) µ m 33 e 41 

In - Ceram 34 - 123 µ m 33 e 41 

Attanasi41 avaliados cor estabilidade do Procera AllCeram dominar e veneering porcelana num estudo in vitro, em condições de envelhecimento acelerado. Depois de 1200 horas de envelhecimento acelerado, a cor mudar entre a pretest e pós - teste de amostras, medido como Δ E valores com o Chroma Meter II instrumento, foi 1,41 ± 1,10 para 300 horas e 1,03 ± 0,62 para 1200 Horas. Nenhuma das amostras Procera demonstrou uma cor mudança que poderia ser detectada clinicamente. 
Medição de dados para atender as marginais do Procera AllCeram coroa não foi relatado. Por conseguinte, o objectivo desta investigação foi a (1) medir os marginais abertura interna e adaptação de Procera AllCeram coroas forjadas para premolar e dentes molares ao trabalho morre, e (2) determinar a precisão de ajuste entre as coroas e os morre atingido por O Procera CAD / CAM sistema para a fabricação de todos - cerâmicos restaurações.

MATERIAL E MÉTODOS 
A parte superior do membro de um Kilgore typodont (Model 500 - B, Kilgore International Incorporated, Coldwater, Mich.), foi usado para fabricar 5 doente modelos de simulação de um arco maxilar. Cinco ivorine maxilar direito firstpremolars e primeiros molares foram preparados com chamfer margens para completar coroas. Dente preparações foram padronizadas com: (1) um total convergência ângulo de 10 graus, (2) chamfer margens 1,3 a 1,5 milímetros circunferencialmente, e (3) oclusal redução de 2,0 milímetros. Todos linha - ângulos foram arredondados. Um premolar e 1 molares foram colocados nas suas posições corretas no maxilar arco do typodont. A plena arco impressão foi feito dos preparativos e dentes adjacentes com policloreto siloxane impressão materiais (Deslocar, Light e Heavy corpo, Kerr Mfg., Romulus, Mich.) em uma bandeja com personalizado verticais pára e, de acordo com instruções do fabricante. Após o período prescrito definição, a impressão foi removido da parte superior do membro do typodont. O preparado ivorine dentes foram removidos do typodont e inserido na impressão, no local certo e orientação. Esta impressão com o preparado ivorine dentes foi derramado em uma melhor dentária pedra para criar um modelo de simulação doente. Cinco pacientes modelos de simulação foram criados como o experimentais população com este método. Por sua vez, uma segunda impressão era feita de cada modelo de simulação para fabricar moldes com mestrado removível premolar e molares morre para cada doente. Final impressões dos modelos de simulação foi derramado em uma resina epoxi materiais (resina epoxi envasamento, General Polymers Corp, Cincinnati, Ohio). Após a resina conjunto, pinos (Pindex System, Whaledent Int. , New York, N.Y.) foram colocados em locais apropriados, bem como a base do elenco foi vertido na cirurgia pedra. Dies foram removidos do elenco base, e talhados à preparação margens. Cinco comandante lança com removível morre foram forjadas para o 5 doente simulações com esta técnica (Fig. 2, Aand B). 
  
Display Full Size versão desta imagem (109K) 

Fig. 2. A, Dies de premolar e dentes molares. B, Cinco conjuntos de matrizes foram forjadas para cada uma das 5 doente modelos de simulação. 
Procera CAD / CAM tecnologia foi utilizada para fabricar todas as restaurações de cerâmica. Neste sistema, a morrer isdigitized por uma varredura instrumento e os dados para o dente preparação são então transmitidos por via electrónica para o laboratório (Procera Sandvik AB, de Estocolmo, Suécia). No laboratório, os dados são utilizados para fabricar os sinterizada aluminous óxido coping.10 Depois de receber a enfrentar desde o laboratório, que é devolvido ao seu respectivo morrem e as características anatômicas do dente são construídos, por parte da cerâmica técnico com Procera AllCeram baixa Espoletas superfície veneering porcelana (Fig. 3, A e B). 
  
Display Full Size versão desta imagem (132K) 

Fig. 3. A, Procera densamente sinterizada cerâmica núcleo para todos - dentes molares. B, Procera AllCeram coroa após baixa fusão superfície porcelana é adicionado ao núcleo para restabelecer a forma anatômica do dente (premolar). 
Laser videografia (Chesapeake Laser System Model LTG - 2105 Lanham, Md.) foi utilizado para medir a precisão de ajuste entre o Procera AllCeram coroas e as matrizes. Esta medição técnica utiliza um laser para digitalizar amostra estrutura e para recolher 3 - dimensionais dados. A Gallium Arsenide laser é a fonte óptica para o sistema que opera no comprimento de onda de 780 mn com a capacidade de medir a micrómetros nível. A precisão deste relatou a laser instrumento é de ± 1,0 µ m 
A medição protocolo usa 3 esferas que funcionam como referências na correspondência a morrer à superfície interna superfície da coroa. Para medir a precisão do ajuste de uma coroa ao morrer, a coroa está posicionado sobre a morrer com uma impressão Silicone materiais como o luting agente (Deslocar Light - corpo, Kerr Mfg. Co.). A fina leigos do silicone é colocado dentro da coroa, que está sentado sobre a morrer e realizada sob uma carga de 1,36 kg para 10 minutos. Quando o Silicone é fixado, a coroa é cuidadosamente removida do morrer, deixando intactos os Silicone com a morrer (Fig. 4). 
  
Display Full Size versão desta imagem (52K) 

Fig. 4. Intacta Silicone morrem abrangendo desde 3 - dimensionais topografia dos aspectos internos do Procera coroa. 
Assim, Silicone prevê: (1) a dimensão física dos 3 - dimensionais espaço entre a coroa e da morte e, (2) a topografia da superfície interna da copa. 
Uma 12,0 x 12,0 milímetros área que totalmente coberto do Silicone material e morrer foi digitalizado usando um x - y - eixo eo eixo medição matriz de 100 µ m entre cada ponto de dados. A superfície da referência esferas também foi digitalizada e seus 3 - dimensionais relações com a morrer foram registrados. Quando digitalização foi concluída, o Silicone luting agente foi removido do morrer e morrem a era digital, sem remoção do digitalizador. A base de dados criada por este processo desde a precisão das medições de aptos para o Procera AllCeram coroas (Fig. 5, A e B). 
  
Display Full Size versão desta imagem (128K) 

Fig. 5. A, Graphic replicação dos aspectos internos do Procera AllCeram coroa. B, Graphic ilustração de digitalização de dados recolhidos pelo laser videografia. 
O sistema a laser videografia medição, através do seu software, é capaz de combinar a superfície do Silicone (superfície interna da copa) à superfície dados da morrem (dente preparação), utilizando o 3 esferas como referências. Uma vez encontradas, o sistema é capaz de selecionar a partir da correspondência coroa - base morrem qualquer número de Secção transversal varredura nas linhas quer o facial ou o proximais direções. Por sua vez, as linhas de varredura são usados para determinar a precisão do ajuste da coroa ao morrer. Desses 2 - dimensionais varreduras, a marginal discrepâncias e lacunas interno pode ser medido em várias localidades entre a morrer e coroa semelhante ao que foi anteriormente reported.42 e 43 Para cada coroa, as seguintes 4 medições locais (Fig. 6), foram Utilizado para determinar a precisão de ajuste entre as coroas e as matrizes para a premolar e os dentes molares: 
1. Marginal Open (PB): A abertura marginais no ponto de aproximação mais estreita entre o epoxi morrem e porcelana margem da coroa. 
2. Axial Wall (AW): A adaptação interna das paredes coroa no ponto médio do axiais parede (2,0 milímetros oclusal à margem da morte). 
3. Cusp Tip (CT): A adaptação interna no cusp dicas. 
4. Oclusal Adaptação (OA): A adaptação interna da superfície oclusal da coroa ao morrer no ponto médio da faciais e proximais. 

  
Display Full Size versão desta imagem (56K) 

Fig. 6. Crown a morrer diagrama mostrando medição localidades ao longo varredura linha utilizada na amostragem de base para determinar marginais abertura, adaptação interna, e precisão de caber na copa de morrer relação. 
O ajuste do Procera AllCeram coroas ao epoxi morre foi avaliado. Ao usar a varredura linha regime previsto para o inquérito, as medições foram tomadas a partir do banco de dados no MO, AW, CT, e OA medição locais para avaliar o ajuste de todas as copas das amostras. A Figura 6 mostra estas medições locais ao longo de uma única linha digitalizar através de um troço de uma coroa de morrer amostra. O premolar exigido 9 digitalizar linhas vista do faciais e proximais superfícies que foram separados por 500 µ m e de 200 µ m, respectivamente. Os molares digitalizar linhas vistas tanto no faciais e as direções proximais foram separados por 700 µ m Os dados gravados nas diferentes linhas de varredura do faciais e proximais varredura linha direções foram calculados pelos diferentes dentes (premolar e molares) e medições locais (MO, AW, CT, e OA). 
Mean dados foram calculados e analisados estatisticamente com estatísticas descritivas e repetida medição analysis44 para avaliar a coroa - entre grupos efeito, dentro de copa grupo medição localidades efeito, e as interações entre grupos e dentro de - coroa - coroa grupos. The Greenhouse - Geisser épsilon (G - G épsilon) ajustamento foi realizado como parte da análise repetida medida com o SuperANOVA program.45 Dados análises para as investigações foram realizadas em .05 nível de significância estatística com o programa SuperANOVA (Abacus Concepts, Berkeley , Na Califórnia). 
A diferença média dimensão em cada medição localização dentro de cada grupo coroa foi determinada pelo cálculo do meio e desvios padrão (DP). Internal adaptação foi determinada pelo cálculo da média e EE da média fosso dimensões internas para a medição localidades por coroa grupo. Precisão de ajuste foi determinada usando a diferença média dimensão de todas as medições locais para calcular a média e SD pela coroa grupo. Repetidas medição análises foram realizadas para a entre - coroa fator (premolar / molares), o fator dentro de copa medição locais (MO, AW, CT, e OA), e as interacções entre estes 2 factores. 
O G - G épsilon é uma probabilidade ajustado para todos os indivíduos dentro - hipótese testes. O epsilons foram concebidos para ajudar a eliminar os efeitos da correlação dentro da medição repetida sobre a validade dos testes. O procedimento utilizado para medir a diferença na medição localidades da coroa - Interface morrer foram considerados medições repetidas dentro de copa grupos experimentais sobre a mesma unidade e foram correlacionados. 
RESULTADOS 
A Figura 7 mostra os meios e EE da média fosso dimensões da coroa medição por grupos locais. 
  
Display Full Size versão desta imagem (38K) 

Fig. 7. Mean fosso dimensão e desvio padrão das copas em grupos marginais abertura (MO), axiais parede (AW), cusp ponta (CT), e oclusal adaptação (OA) medição locais para premolar e molares coroas. 
O fosso PB dimensão média e SD para o premolar e molares coroas foram 56,0 µ m SD ± 21 µ m e 63,0 µ m SD ± 13 µ m, respectivamente. A adaptação interna fosso dimensões meios e EE para a premolar coroas foram 69,0 µ m SD ± 17 µ m (AW), 48,0 µ m SD ± 12 µ m (CT), e 36,0 µ m SD ± 7 µ m (OA). Embora a adaptação interna fosso dimensões meios e EE para a molares coroas foram 49,0 µ m SD ± 3 µ m (AW), 67,0 µ m SD ± 21 µ m (CT), e 74,0 µ m SD ± 29 µ m (OA) , A precisão de caber fosso dimensão meios e EE para a premolar e molarcrowns foram 52,0 µ m SD ± 19 µ m e 63 µ m SD ± 20 µ m, respectivamente (Fig. 8). 
  
Display Full Size versão desta imagem (51K) 

Fig. 8. Mean fosso dimensão e desvio padrão das copas a morrer precisão dos próprios para premolar e molares coroas. 
A precisão do ajuste fosso dimensões para a coroa grupos foram iguais ou inferiores a coroa grupos PB diferença média dimensão. 
Quadro II mostra a análise de variância (ANOVA) da medição análises repetidas. 
Quadro II. 
Medidas repetidas análise por entre copas - Factor (premolar / molares), dentro de medição fator localização (MO, AW, CT, e OA), bem como a interacção dos 2 factores 
Fonte df Soma de quadrados Mean praças valor F G - G valor 
Crown 1 0,0012 0,0012 1,8934 
Assunto (grupo) 8 0,0052 0,0007 
Medição localização 3 0,0001 4.6733E - 5 0,2647 0,8098 
Medição localização * copas 3 0,0044 0,0015 8,3687 0,0015 * 
Medição localização * sujeito (coroa) 24 0,0042 0,0002 
G - G Epsilon 
ÅMeasurement localização 0,8082 
* Significativos em .05 nível. 

Os valores reportados P foram modificadas pelo G - G épsilon. Os fatores épsilon (å), para o grau de liberdade (df) ajustamento são fornecidos no canto inferior esquerdo. As probabilidades não são corrigidos para valores de épsilon superior a 1. 

Não se observaram diferenças significativas encontradas entre as copas grupos fosso dimensões ou dentro coroa medição grupos locais fosso dimensões (G - G = 0,8098). As interacções do 2 fatores foram significativos (G - G = 0,0015). 
Para identificar a fonte de interacções significativas, os dados foram analisados mais. Devido a repetidas medida análise contida dentro e entre os fatores, o emparelhados comparações não serão permitidas. Portanto 1 - forma ANOVA46 sobre a copa foi realizada. Quadro III apresenta a 1 - forma ANOVAs da coroa grupos por MO, AW, CT, e OA medição locations.Mean fosso entre as dimensões coroa grupos foram significativamente diferentes em AW e OA medição locais. MO e CT medição localidades diferença média dimensões, não demonstram qualquer diferença significativa entre os grupos coroa. Para realizar comparação análise da medição dentro copas locais fosso dimensões, os dados foram separados por coroa grupo e analisados com 1 - dentro fator medidas repetidas ANOVA pela coroa fator (premolar / molares). 
Quadro III. 
- Uma maneira ANOVAs de entre - coroa fator de medição locais (MO, AW, CT, e OA) 
Fonte df Soma de quadrados Mean praças valor F P valor 
MO Crown 1 0,0001 0,0001 0,3812 .5541 
Residual 8 0,0024 0,0003 
AW Crown 1 0,0010 0,0010 6,6000 .0332 * 
Residual 8 0,0012 0,0001 
CT Crown 1 0,0009 0,0009 3,0669 .1180 
Residual 8 0,0024 0,0003 
OA Crown 1 0,0036 0,0036 8,5186 .0193 * 
Residual 8 0,0034 0,0004 

Diferença significativa em .05 nível. 
A ANOVA dos repetidos medida análise para o premolar coroas dentro de medição fator localização (Quadro IV) revelaram diferenças significativas na medição localidades fosso dimensões (G - G = 0,0293), enquanto que os molares coroas mostrada nenhuma diferença significativa na medição Localidades fosso dimensões (Tabela V). No entanto, a medição localização fosso dimensões foram aproximando um marginais nível de diferenças significativas (G - G = 0,1423). As interacções entre a medição localização e coroa não eram significativos, quer para coroar grupo. 
Quadro IV. 
- Dentro de um fator (premolar) repetido medidas ANOVA sobre a medição fator localização (MO, AW, CT, e OA). 
Fonte df Soma de quadrados Mean praças valor F G - G valor 
Premolar 4 0,0016 0,0004 
Medição localização 3 0,0030 0,0010 5,7164 0,0293 * 
Medição localização * premolar 12 0,0021 0,0015 
G - G Epsilon 
ÅMeasurement localização 0,6591 


Significativos em .05 nível. 
P relatados valor foi modificada pela G - G épsilon. Os fatores épsilon (å), para o grau de liberdade (df) ajustamento são fornecidos no canto inferior esquerdo. As probabilidades não são corrigidos para valores de épsilon superior a 1. 
Tabela V. 
- Dentro de um fator (molares) repetido medidas ANOVA sobre medição fator localização (MO, AW, CT, e OA) 
Fonte df Soma de quadrados Mean praças valor F G - G valor 
Molares 4 0,0036 0,0009 
Medição localização 3 0,0016 0,0005 2,9492 0,1423 
Medição localização * molares 12 0,0021 0,0002 
G - G Epsilon 
ÅMeasurement localização 0,4406 

P relatados valor foi modificada pela G - G épsilon. Epsilon factores (å), para o grau de liberdade (df) ajustamento são fornecidos no canto inferior esquerdo. Probabilidades não são corrigidos para valores de épsilon superior a 1. 
Para comparar os locais dentro coroa medição fosso dimensões, o contraste test47 foi realizada no .05 nível de significância. Três medições locais apresentaram diferenças significativas em fosso dimensões para o premolar coroas (Tabela VI). Os molares coroas tinham 1 medição local com um gapdimension de diferença significativa (G - G = 0,0423). A AW e OA fosso dimensões foram aproximando marginais significância (G - G = 0,0960) na molares coroas. 
Quadro VI. 
Contraste comparação teste de copa fator (premolar / molares), para o efeito pela medição fator localização (MO, AW, CT, e OA) 
Fonte df Soma de quadrados Mean praças valor F G - G valor 
Premolar PB versus AW 1 0,0004 0,0004 2,5721 0,1466 
X CT 1 0,0002 0,0002 0,9632 0,3112 
Tm OA 1 0,0010 0,0010 5,8450 0,0512 * 
AW versus CT 1 0,0012 0,0012 6,6833 0,0411 * 
Tm OA 1 0,0028 0,0028 16,6833 0,0065 * 
CT versus OA 1 0,0004 0,0004 2,0627 0,1805 
Molares PB versus AW 1 0,0004 0,0004 2,4384 0,1563 
X CT 1 4.4100E - 5 4.4100E - 5 0,2469 0,4481 
Tm OA 1 0,0003 0,0003 1,7555 0,1971 
AW versus CT 1 0,0008 0,0008 4,2370 0,0960 
Tm OA 1 0,0015 0,0015 8,3318 0,0432 * 
CT versus OA 1 0,0001 0,0001 0,6857 0,3210 
* Significativos em .05 nível. 

Os valores reportados P foram modificadas pelo G - G épsilon. A célula pesos utilizados no contraste teste eram de mais (+) 1,0000 a menos (-) 1,0000. 
DISCUSSÃO 
An PB aceitável para o pleno coroas, conforme relatado por Hung e Al. , 26 é 50 a 75 µ m, enquanto Weaver e al.29 sugeriram 70 µ m SD ± 10 µ m A diferença média MO dimensões para o premolar e molares coroas grupos desta investigação foram 56,0 µ m SD ± 21 µ m e 63,0 µ m SD ± 13 µ m, respectivamente. Muitos disponíveis todos - cerâmicos coroa sistemas demonstrar muito maior fosso PB dimensões (Quadro I). Os MOs de todas as copas - cerâmicos sistema testados neste estudo são comparáveis e, em alguns casos, melhor do que o fosso dimensões reportadas por outros systems.17, 20, 21, 25, 26, 27, 28, 29, 30, 31 , 32, 33, 34, 35, 36, 37, 38, 40 e 41 Mean MOs entre a coroa grupos no presente inquérito não foram significativamente diferentes. 
A abertura absoluta marginais (SAM), conforme descrito por Holmes e Al. , 42 pode ser demonstrada quando uma linha conecta a "A e E" pontos de referência na Figura 6. A AMO é relatado na maioria dos estudos da coroa caber usando MO como um termo genérico. A AMO não foi utilizado neste estudo, uma vez que foi o consenso dos investigadores que a AMO iria variar muito, dependendo da overextension ou underextension da coroa amostra. A MO, a abertura demonstrada quando uma linha conecta a "B e E" pontos de referência, não será afectada pela overextension ou underextension das coroas (Fig. 6). Clinicamente, é importante conhecer bem como uma coroa encaixar o dente na margem, mas igualmente importante é a dimensão do luting space.13 O cimento espaço interno ou adaptação é considerada um espaço uniforme que facilita lugares sem comprometer a retenção ea Resistência formas. Isto é de extrema importância, pois todos - cerâmicos coroas são mais frágeis do que o elenco ligas alternativas. A precisão do ajuste da coroa podem influenciar a clínica prognosis.Tuntiprawon e Wilson13 informou que todos - cerâmicos coroas exibido compressão maior força quando a média AW estava em um fosso dimensão de 73,0 µ m Their estudo também mostrou que, quando a média AW foi aumentado para 122,0 µ m, uma menor resistência fracasso ocorreu sem qualquer melhoria significativa nos assentos. Por conseguinte, um conjunto de cerâmica copa que pode ser previsivelmente fabricados com a precisão de ajuste e não comprometa o dente tem o melhor prognóstico para o sucesso clínico. Os resultados do estudo mostram que as coroas estudados serão sempre lugar, com uma diferença PB dimensão de menor intensidade do que outros, mostraram estar associados à óptima todos - cerâmicos coroa strength.13 A diferença dimensão aumento de 49 µ m acima dos ideais diferença (73,0 µ m ), Para o AW dimensão só diminuiu a diferença média dimensão PB aproximadamente 3,0 µ m.13 Portanto aceitável uma diferença AW dimensão deve estar dentro de 1 SD do ideal AW dimensão. A média dimensão AW diferença para o premolar coroas (69,0 µ m SD ± 17 µ m) não foi significativamente diferente da literatura relatório, mas o AW para os molares foi 23,0 µ m menos (49,0 µ m ± 3 µ m SD) . Os dados não indicam que houve incidências axiais de contato entre a parede e as coroas morre porque o contacto entre as paredes axiais causaria impingement e afectar lugares sentados. No entanto, a diferença média dimensão em todas as medições locais foram menos do que a média fosso dimensões mostraram - se estar associada com maior incumprimento força da porcelana casaco crown.13 Assim, quando precisão de ajuste é combinado com strength11, 12 e 13 e cor estabilidade, 42 o prognóstico clínico é reforçada. 
A 1 - forma ANOVAs (Quadro III) revelou que a média dimensões fosso entre copas grupos foram significativamente diferentes para AW e OA medição locais. Figura 7 ilustra a relação inversa a medição localidades significa ter dimensões fosso entre copas grupos. O premolar claramente tem uma dimensão maior do que o fosso molares AW medição no local. Contudo, não existem diferenças significativas nas dimensões lacuna no CT medição coroa entre grupos locais, mas a dimensão molares diferença foi superior ao premolar. Gap dimensões na OA medição localidades entre copas grupos foram significativamente diferentes, e os molares dimensão diferença foi superior ao premolar. Premolar fosso diminuiu de dimensões AW> CT> OA, enquanto que os molares fosso dimensões aumentaram de AW <CT <OA entre copas grupos. 
A luting espaço de 49 µ m no AW para os molares era muito menor do que os 69 µ m para o premolar. A maior diferença AW dimensão no premolar grupo combinado com uma pequena diferença PB dimensão sugerem que outros assentos poderia eventualmente ocorrer. A maior diferença dimensões no CT (67 µ m) e OA (74 µ m) localizações para os molares, quando comparados com os premolars CT (48 µ m) e OA (36 µ m) também sugerem incompleto lugares sentados. No entanto, CT média dimensão fosso entre copas grupos não suporta incompleto lugares sentados. Estatisticamente, um 19 µ m diferença entre a coroa grupos CT fosso dimensões não foi significativa, enquanto um 38 µ m diferença na OA foi significativa. Teoricamente, uma diminuição de 19 µ m, no CT no molares grupo continuaria a deixar uma lacuna dimensão de 19 µ m na OA. Uma melhoria desta magnitude no fosso dimensão para a PB e AW medição locais, não seria viável. A análise mostra um grupo diferença entre copas em média MOs de 7,0 µ m Se uma lacuna redução de 19 µ m ocorreu no CT como anteriormente afirmado, a diferença de dimensões fosso entre a coroa grupos MOs não deve alterar a inferência estatística. Isto poderá levar a uma pergunta se as diferenças na AW, CT, e OA eram verdadeiras diferenças no fabrico das coroas, ou seja, as diferenças foram causados pelo método de medição. Na melhor das hipóteses, a diferença global da coroa grupos MOs seria 12,0 µ m com a hipotética 19,0 µ m melhoria. Gap dimensões na MO e CT medição localidades foram comparáveis entre os grupos coroa. Se houve erros nos métodos utilizados para medir a caber, em seguida, o banco de dados será composta por valores inflacionados e do ajuste das coroas seria muito melhor do que os nossos resultados relatados. A mais importante questão a resolver é a de saber se as diferenças são significativas clinicamente. Obviamente, o AW OA e podem influenciar a pressão hidráulica criada durante a cimentação procedimentos independentemente da viscosidade da mistura de cimento. If the luting agent and method used in this study was the cause of the significant differences recorded in the gap dimensions, the fit of the crowns is still an improvement in technology. The MO mean gap dimensions for crown groups were comparable to previous reports.18, 21, 22, 26, 27, 28, 29, 30, 31, 32, 33, 34, 35, 36, 37, 38, 39, 41 and 42 The internal adaptation and precision of fit for the crowns were comparable to the MO gap dimensions. 
The 3 most common instruments used to measure the fit of the crowns to the supporting structure are stereomicroscopy, traveling microscopy, and videomicroscopy. The method used with most of these instruments provide limited databases generated from widely separated data points. The measurements recorded with these instruments are taken at a single location per axial surface at the margin of the crown. The mean dimension calculated for a specific measurement location, for example, the MO, have often demonstrated excessive standard deviations. Subsequently, questions have been raised regarding the repeatability of these methods and the adequacy of the database used in the documentation of the crowns fit. Chan et al.24 and Karlsson25 reported differences in the MOs between a crown and die in the range of 54 to 130 μm within a distance of 300 μm along margins of the restorations. The laser method used in our study generated an average of 8,000 and 12,000 data points for the premolar and molar dies, respectively. From this database, 9 scan lines in the facial and proximal directions were selected for each matched crown to die sample. From the 18 scan lines, 36 measurements were made for each MO, AW, and CT measurement location, whereas 18 measurements were recorded for the OA measurement location for a total of 126 measurements per crown and die sample. Thus, a total of 252 points per sample were used to analyze the fit of each crown. The repeated measures ANOVA showed no significant differences in the average gap dimension for the 4 measurement locations between the crown groups, and within crown group measurement locations gap dimensions had no significant effect. The significant interaction (G-G=0.0015) of the measurement location gap dimension by crown is of interest because it indicates that the differences in measurement location gap dimension is dependent on the crown group and the location. 
The separate within-factor repeated measures ANOVAs indicated that the premolar measurement location gap dimensions were significantly different (G-G=0.0293, Table IV), whereas the molars were not significantly different (G-G=0.1423, Table V ). Although the differences in the measurement locations gap dimensions for the molars were not significant at the level of .05, maybe a significant difference would occur with a larger sample size or with a cement of a different viscosity (Table V). 
To isolate the significant differences in gap dimensions within crown groups, separate contrast comparison tests were performed by the measurement locations. The within-crown group differences in gap dimensions were clearly demonstrated (Table VI). The smaller crowns had significant differences in measurement locations gap dimensions when comparing the MO to OA, the AW to CT, and the AW to OA, whereas the larger crowns demonstrated significant differences only in the AW to OA measurement locations gap dimensions. These findings further support the position that the differences in the gap dimensions are dependent on the crown and the measurement location. 
Another possible explanation for the difference within crown groups could be related to the refractory die fabrication process. The difference in the calculation and the ability of the Procera CAD/CAM to perform at the level of accuracy required to fabricate the refractory die in the suggested 15% to 20% enlargement range9 and 10 may be the cause of these differences. This is supported by the findings that there were more significant differences in the gap dimensions of the measurement locations of the crowns with the smallest surface area. This requires further study, beyond the scope of this investigation, to determine the cause of these significant differences. 
Mean gap dimensions for the MO measurement location in Figure 1 is of particular interest to clinicians because it represents the gap between the external surface of the tooth and the crown margins. It is paramount that the MO be kept to a minimum to establish clinical acceptability of the fit for any crown system. Ideally this space should be in the 25 to 40 μm range to accommodate the different types of luting agents, but most often the MOs are greater. In our study, the luting space was recorded by positioning the crowns on the respective dies with a light-bodied silicone to simulate the luting agent. This approach has been used by other investigators in measuring the gap dimension.27, 37 and 38 McLean and von Fraunhofer38 demonstrated that cementation simulation with polyether rubber will require a space of 22 μm, whereas 20 μm is needed for zinc-phosphate cement. Therefore it is theorized that the luting space measured and represented by the silicone impression did not prevent the accurate seating of the crown as a result of hydraulic pressure. Furthermore, data analysis confirmed that no impingements occurred at the measurement locations used. The available space is more than enough to accommodate the space required for the commonly used Type I and II luting agents of approximately 25 to 40 μm. 
CONCLUSIONS 
This study used an in vitro model and ivorine teeth and a light body impression material functioning in a dual role to simulate cement and to capture the accurate topography of the internal aspect of the crowns. The following conclusions were drawn. 
1. Quantitatively generated database of the precision of fit at the crown-die interface using laser videography combined with the provisional cementation technique appears to be a valid method to quantify the marginal gap, cement space, or to determine the precision of fit at the crown-die interface. 
2. The Procera AllCeram crown for the posterior dentition demonstrated a marginal gap at the crown-die interface that was less than 70 μm. 
3. Precision of fit at the crown-die interface for premolar and molar crowns was 56 ± 21 μm and 63 ± 13 μm, respectively. 
4. The difference in the fit between crown groups depend on the region of the crown being evaluated and the type of crown. 
5. The Procera CAD/CAM system consistently produced all-ceramic crowns for premolar and molar teeth with a clinically acceptable fit at the margin (54 to 64 μm) and internal aspects (49 to 63 μm), which is conducive to clinical success. 

Acknowledgements 
We would like to acknowledge the assistance of Beth E. Lang in the operation of the laser digitizer and collection of data. We also thank Mr. Rui-Feng Wang for his assistance in the preparation of the measurement data and the performance of the several statistical tests used in this investigation 

References 
1 Cerestore crown technical manual. Fabrication process. Lakewood (CO): Coors Biomedical Co. 
2 Dicor laboratory technical manual, Dentsply International Inc, York (1987). 
3 KH. Koerber, Ceraplatin crown structures with reproducible minimal thickness strength, Huethig, Heidelberg (1985). 
4 JG Wall and DL. Cipra, Alternative crown systems. Is the metal-ceramic crown always the restoration of choice?, Dent Clin North Am 36 (1992), pp. 765–782. View Record in Scopus | Cited By in Scopus (12) 
5 JW McLean and TH. Hughes, The reinforcement of dental porcelain with ceramic oxides, Br Dent J 119 (1965), pp. 251–267. 
6 S. Timoshenio, Theory of plates and shells, McGraw-Hill, New York (1959) 79. 
7 SF Hulbert, JC Brokros, LL Hench, J Wilson and G. Heimke, Ceramics in clinical applications, past, present and future. In: P Vincenzinc, Editor, High tech ceramics., Elsevier Science, Amsterdam (1987), pp. 3–27. 
8 P. Boutin, Total arthroplasty of the hip by fitted aluminum prosthesis. Experimental study and 1st clinical application, Rev Chir Orthop Appar Mot 58 (1972), pp. 229–246 [in French]. View Record in Scopus | Cited By in Scopus (99) 
9 M Andersson and A. OdÉn, A new all-ceramic crown: a dense-sintered, high-purity alumina coping with porcelain, Acta Odontol Scand 51 (1993), pp. 59–64. View Record in Scopus | Cited By in Scopus (90) 
10 M Andersson, B Bergman, C Bessing, G Ericson, P Lundquist and H. Nilsson, Clinical results with titanium crowns fabricated with machine duplication and spark erosion, Acta Odontol Scand 47 (1989), pp. 279–286. View Record in Scopus | Cited By in Scopus (57) 
11 WC Wagner, TM Chu and ME. Razzoog, Indentation fracture toughness of three new ceramic core materials, J Dent Res 74 (1995), p. 471 (Abstract 568). Abstract | PDF (1118 K) | View Record in Scopus | Cited By in Scopus (11) 
12 WC Wagner and TM. Chu, Biaxial flexural strength and indentation fracture toughness of three dental core ceramics, J Prosthet Dent 76 (1996), pp. 140–144. SummaryPlus | Full Text + Links | PDF (1212 K) | View Record in Scopus | Cited By in Scopus (69) 
13 M Tuntiprawon and PR. Wilson, The effect of cement thickness on the fracture strength of all-ceramic crowns, Aust Dent J 40 (1995), pp. 17–21. View Record in Scopus | Cited By in Scopus (17) 
14 American Dental Association, ANSI/ADA Specification No. 8 for zinc phosphate cement. In: Guide to dental materials and devices (5th ed), American Dental Association, Chicago (1970–71). 
15 PJ Adair and DG. Grossmann, The castable ceramic crown, Int J Periodont Restor Dent 4 (1984), pp. 33–45. 
16 C Chan, G Haraszthy, J Geis-Gerstorfer and H. Weber, The marginal fit of Cerestore full-ceramic crowns—a preliminary report, Quintessence Int 6 (1985), pp. 399–402. 
17 DG. Grossman, Cast glass ceramics, Dent Clin North Am 29 (1985), pp. 725–739. View Record in Scopus | Cited By in Scopus (21) 
18 T Sato, A Wohlend and P. Schärer, “Shrink-free” ceramic crown system: factors influencing the core marginal fit, Quintessence Dent Technol 10 (1986), pp. 81–86. 
19 CE Brukl and GK. Philp, The fit of molded all-ceramic, twin foil, and conventional ceramic crowns, J Prosthet Dent 58 (1987), pp. 408–413. Abstract | View Record in Scopus | Cited By in Scopus (3) 
20 DR. Davis, Comparison of fit of two types of all-ceramic crowns, J Prosthet Dent 59 (1988), pp. 12–26. 
21 P Schärer and T. Sato, A comparison of the marginal fit of three cast ceramic crown systems, J Prosthet Dent 59 (1988), pp. 534–542. 
22 A Wohlwend, JR Strub and P. Schärer, Metal ceramic and all-porcelain restorations: current considerations, Int J Prosthodont 2 (1989), pp. 13–26. View Record in Scopus | Cited By in Scopus (17) 
23 C Chan, G Haraszthy, J Geis-Gerstorfer, H Weber and H. Huettemann, Scanning electron microscopic studies of the marginal fit of three esthetic crowns, Quintessence Int 20 (1989), pp. 189–193. View Record in Scopus | Cited By in Scopus (9) 
24 S. Karlsson, The fit of Procera titanium crowns. An in vitro and clinical study, Acta Odontol Scand 51 (1993), pp. 129–134. View Record in Scopus | Cited By in Scopus (31) 
25 MF Abbate, AHL Tjan and WM. Fox, Comparison of the marginal fit of various ceramic crown systems, J Prosthet Dent 61 (1989), pp. 527–531. Abstract | View Record in Scopus | Cited By in Scopus (29) 
26 SH Hung, KS Hung, JD Eick and RP. Chappell, Marginal fit of porcelain-fused-to-metal and two types of ceramic crowns, J Prosthet Dent 63 (1990), pp. 26–31. Abstract | View Record in Scopus | Cited By in Scopus (12) 
27 F Vahidi, ET Egloff and FV. Panno, Evaluation of marginal adaptation of all-ceramic crowns and metal ceramic crowns, J Prosthet Dent 66 (1991), pp. 426–431. Abstract | View Record in Scopus | Cited By in Scopus (3) 
28 KM Harrison, EJ Billy and GB. Pelleu, Effects of an internal escape channel on a castable ceramic crown, J Prosthet Dent 65 (1991), pp. 622–627. Abstract | View Record in Scopus | Cited By in Scopus (0) 
29 JD Weaver, GH Johnson and DJ. Bales, Marginal adaptation of castable ceramic crowns, J Prosthet Dent 66 (1991), pp. 747–753. Abstract | View Record in Scopus | Cited By in Scopus (13) 
30 T Hummert, N Barghi and T. Berry, Post cementation marginal fit of a new ceramic foil crown system, J Prosthet Dent 68 (1992), pp. 766–770. Abstract | View Record in Scopus | Cited By in Scopus (2) 
31 A Lomanto and S. Weiner, A comparative study of ceramic crown margins constructed using different techniques, J Prosthet Dent 67 (1992), pp. 773–777. Abstract | View Record in Scopus | Cited By in Scopus (0) 
32 RJ Hoard, PC Chiang, ER Hewlett and AA. Caputo, Marginal discrepancy as related to margin design in porcelain-fused-to-Dicor restorations, Oral Health 83 (1993), pp. 15–16. 
33 NJ Grey, V Piddock and MA. Wilson, In-vitro comparison of conventional crowns and a new all-ceramic system, J Dent 21 (1993), pp. 47–51. Abstract | View Record in Scopus | Cited By in Scopus (20) 
34 AJ Schonenberger, A Di Felice and M. Cossu, Improving the precision of esthetic ceramic margins: guidelines for success, J Esthet Dent 6 (1994), pp. 143–150. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (0) 
35 P Pera, S Gilodi, F Bassi and S. Carossa, In vitro marginal adaptation of alumina porcelain ceramic crowns, J Prosthet Dent 72 (1994), pp. 585–590. SummaryPlus | Full Text + Links | PDF (2114 K) | View Record in Scopus | Cited By in Scopus (18) 
36 MI Fanuscu and JA. Sorensen, Reduced microleakage of adhesive ceramic crowns by modified cementation methods. IADR, J Dent Res 72 (1993), p. 186 (Abstract 663) (Special Issue). 
37 JW McLean and JA. von Fraunhofer, The estimation of cement film thickness by an in vivo technique, Br Dent J 131 (1971), pp. 107–111. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (68) 
38 JW. McLean, Polycarboxylate cements. Five years' experience in general practice, Br Dent J 132 (1972), pp. 9–15. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (1) 
39 AJ Dickinson, BK Moore, RK Harris and RW. Dykema, A comparative study of the strength of aluminous porcelain and all-ceramic crowns, J Prosthet Dent 61 (1989), pp. 297–304. Abstract | View Record in Scopus | Cited By in Scopus (13) 
40 S Rinke, G Margraf, L Jahn and A. Huls, The quality appraisal of copy-milling complete-ceramic crown structures (Celay/In-Ceram), Schweiz Monatsschr Zahnmed 104 (1994), pp. 1495–1499 [in German]. View Record in Scopus | Cited By in Scopus (5) 
41 RC. Attanasi Jr, Color stability of Procera all-ceramic porcelain. [Masters thesis.]. In: H Horace, Editor, Rackham School of Graduate Studies., University of Michigan, Ann Arbor (1995). 
42 S Valderrama, N Van Roekel, M Andersson, CJ Goodacre and CA. Munoz, A comparison of the marginal and internal adaptation of titanium and gold-platinum-palladium metal ceramic crowns, Int J Prosthodont 8 (1995), pp. 29–37. View Record in Scopus | Cited By in Scopus (19) 
43 JR Holmes, SC Bayne, GA Holland and WD. Sulik, Considerations in measurement of marginal fit, J Prosthet Dent 62 (1989), pp. 405–408. Abstract | View Record in Scopus | Cited By in Scopus (42) 
44 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 1049-57. 
45 BS. Everitt, The Cambridge dictionary of statistics in the medical sciences, Cambridge, Melbourne (1995) 111. 
46 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 529-32. 
47 J Neter, W Wasserman and MH. Kutner, Applied linear statistical models (3rd ed), Irwin, Boston (1990) 576-77.