Material characteristics of a novel shrinkage-free ZrSiO4 ceramic for the fabrication of posterior crowns

G. Heydeckea, b, , , F. Butza, J.R. Binderc and J.R. Struba 
aDepartment of Prosthodontics, University Hospital, Freiburg, Germany
bFaculty of Dentistry, McGill University, Montreal, Canada
cInstitute for Materials Research III, Forschungszentrum Karlsruhe, Karlsruhe, Germany 
Received 15 August 2005;  revised 21 March 2006;  accepted 22 June 2006.  Available online 1 September 2006.

- Abstract

- Objective

To confirm the clinical applicability of a novel ZrSiO4 (zircon) based shrinkage-free ceramic material, the flexural strength, fracture toughness and chemical solubility were tested. In addition, the fracture load of full crowns made from this material was tested after cyclic thermomechanical loading.

- Methods

Flexural strength of 12 specimens was measured using a biaxial bending test. Fracture toughness was measured using 10 slotted box shaped specimens. The specimens were fractured using a universal testing machine; fracture loads were recorded. A chemical solubility test was performed in accordance with ISO norm 6872. Additionally, 32 ZrSiO4 all-ceramic crowns were fabricated on extracted caries-free human molars. Sixteen Empress 2 and 16 PFM crowns served as a reference control. After artificial aging of 1.2 million cycles in the chewing simulator, the survival rate of the crowns was determined. The fracture load of all surviving specimens was obtained by loading the crowns until fracture in a universal testing machine.

- Results

A flexural strength of 328 MPa was found. The fracture toughness of the ZrSiO4 ceramic was 5.16 MPa m0.5. The chemical solubility amounted to 7.2 μg/cm2. All specimens survived the chewing simulation (survival rate: 100%); no crowns had to be re-cemented. A mean fracture strength of 1790 N was found for Everest HPC for Empress 2 crowns, 1715 N for Empress 2 crowns and 2416 N for PFM crowns. Fracture loads of PFM crowns were significantly higher than for Empress 2 crowns (P = 0.032) as well as ZrSiO4-crowns (P = 0.007). There was no significant difference between ZrSiO4-crowns and Empress 2 crowns (P = 0.743).

- Significance

At the present stage, Everest HPC can be recommended for the fabrication of single crowns as an alternative to conventional PFM and other all-ceramic crowns, because its fracture strength exceeds average masticatory forces in the posterior region.

Keywords: All-ceramic; Fracture strength; Crowns; Fatigue loading

- Article Outline

1. Introduction

2. Materials and methods

2.1. Material tests

2.1.1. Flexural strength

2.1.2. Fracture toughness

2.1.3. Chemical solubility

2.2. Testing of CAD/CAM machined crowns

2.2.1. ZrSiO4-crowns

2.2.2. Porcelain-fused-to-metal crowns

2.2.3. IPS Empress 2 crowns

2.2.4. Luting procedure

2.2.5. Fatigue tests: thermomechanical loading

3. Results

3.1. Material tests

3.1.1. Flexural strength

3.1.2. Fracture toughness

3.1.3. Chemical solubility

3.2. Testing of CAD/CAM machined crowns

4. Discussion

References

- 1. Introduction

Masticatory efficiency, phonetics and esthetic appearance are essential determinants of patient satisfaction with dental prostheses [1]. However, the individual evaluation of esthetics is very subjective [41]. Over the past decades, patient expectations of dental esthetics have been increasing continuously, particularly as they relate to the selection of tooth color [10]. A high degree of sophistication has been achieved in the realm of tooth color esthetics. Three factors can be held responsible for this development: (1) the influence of highly skilled dental technicians, (2) the continuous refinement of manufacturing techniques, and (3) improvement of dental materials [3], [11] and [14]. However, esthetic alternatives to the full gold crown such as porcelain-fused-to-metal (PFM) crowns are generally more expensive. All-ceramic crowns are even more resource intensive with regard to their manufacture and cost. Social insurance systems mostly cover the least costly option – making the cast gold crown the standard of care.

In addition to the PFM technique, there are quite a number of different systems available for the manufacture of all-ceramic restorations. However, all of those systems were introduced with the aim of further improving the appearance of tooth colored fixed prosthodontics and to extend their indication to the posterior area. These can be considered top end restorations. There are few efforts to improve esthetics at the standard level. CAD/CAM techniques appear to the most suitable for the creation of cost effective fixed restorations. But, so far, they have not been explored to their full potential to provide simple and low cost single crowns. If asking patients, most will answer that natural teeth are “white”. While this may seem a coarse definition to a dental professional this could be a starting point for simple alternatives to a cast gold crown which is not white.

Several techniques have been described for the fabrication of all-ceramic fixed restorations. For example, with aluminium oxide glass infiltrated core ceramics (In-Ceram alumina, Vita, Germany) a flexural strength of >450 MPa can be achieved. With heat pressed lithium–disilicate cores, flexural strengths of slightly less than 400 MPa have been reported [36]. Both types of materials required highly skilled manual techniques for their respective slip-casting, waxing or heat pressing procedures. One of the drawbacks of all sintered ceramics is their linear sinter shrinkage of up to 20%. To obtain the desired dimensions, restorations must be formed to correspondingly enlarged dimensions.

More recent fabrication techniques involve the machining of CAD-designed restorations (e.g., Cerec, Sirona, Germany; DCS, Switzerland; Cercon; Degudent, Germany). The most prevalent core material is Y-TZP zirconia. The machining of restorations from sintered ZrO2 is quite time consuming; the machining process for a single unit restoration often takes 2–3 h [5]. In addition, the extreme hardness of sintered high strength ceramics results in high wear of the machining tools. Recent evidence shows that the machining process also has a detrimental effect on the microstructure of the material. As a consequence, the induced microcracks reduce the reliability of restorations machined from fully sintered ZrO2 [28] and [29]. Restorations made using these techniques mostly require a veneering layer to (a) achieve a proper tooth shape and (b) to cover the core material to achieve acceptable esthetics.

An alternative to slip-casting or grinding of sintered blocks is CAD/CAM machining of pre-sintered blocks or blanks at the green stage. The use of a conventional ceramic material process requires an enlarged green body to compensate for the linear sinter shrinkage. The enlargement factor depends on the density of the green compact, the density of the body after sintering and their respective weights. Thus, the precision of the white body directly depends on a constant quality of the green bodies and the stability of the sintering process.

These problem areas can be circumvented if a shrinkage-free ceramic material is used. With a so called reaction bonding process, this is achieved through the volume expansion of one component during the sintering process compensating the volume loss of another. If a low loss binder is added, the sinter shrinkage can be reduced further. Such a system based on zirconium silicide (ZrSi2), zirconia (ZrO2) and a polymethylsilsesquioxane (PMSS) has been described in detail elsewhere [17]. The ceramic based on this system is zircon (ZrSiO4). It can be molded to blanks using axial or isostatic pressing. The blanks are stable and can be machined using commercial CAD/CAM equipment.

The aims of this in vitro study were as follows:

1. to characterize the zircon (ZrSiO4) further with testing procedures relevant to dental ceramics;

2. to assess the survival rate of all-ceramic crowns made from zircon ceramic after 1.2 million load cycles in a dual axis chewing simulator;

3. to measure the load to fracture of all-ceramic crowns made from zircon ceramic after 1.2 million load cycles in a dual axis chewing simulator.

- 2. Materials and methods

To test the new ZrSiO4 ceramic (Everest HPC, KaVo, Biberach, Germany) regarding its suitability for dental restorations, a number of tests were carried out. This included flexural strength testing, calculation of the Weibull modulus, the measurement of the fracture toughness, as well as the chemical solubility.

- 2.1. Material tests

- 2.1.1. Flexural strength

Flexural strength was measured using a biaxial piston-on-3-ball test. Twelve disk shaped specimens were prepared with a diameter of 15.1 ± 0.1 mm and a thickness of 1.2 ± 0.1 mm. The specimens were measured with a micrometer screw and their actual measures recorded. Then, the specimens were placed in the testing device and testing was then carried out in a universal testing machine (Zwick, Germany) at a crosshead speed of 1.5 mm/min. The flexural strength σ was calculated using the formula:

where ; υ = 0.19 (Poisson's parameter for ceramic); r1 = 5.97 mm (radius of the circle of the supporting balls); r2 = 0.76 mm (radius of the piston); r3 = radius of the specimen in mm; Fmax = fracture load in N; d = thickness of the specimen in mm.

In addition, the Weibull strength σ0 and the Weibull modulus m were calculated using linear regression analysis. The fracture probability Pσ is defined by the following equation: Pσ = (n − 0.5)/N, where N is the total number of specimen and n refers to the ascending order of the σ-values. The natural logarithms of σ are calculated and entered as x-values. The corresponding y-values are calculated using the formula ln (ln (1/(1 − Pσ))) and are then inserted into the coordinate plane. The diagram shows a straight line with the slope m and the intersection with the x-axis ln σ0.

- 2.1.2. Fracture toughness

Fracture toughness was measured using 10 box shaped specimens measuring 20 mm × 3 mm × 1.5 mm. Before testing, all measures of the specimens were recorded. Then, a notch was prepared within each specimen to a depth of 50% using a diamond jig saw. The width of the notch was defined by the thickness of the saw blade of approximately 300 μm. The specimens were then fractured in the universal testing machine at a crosshead speed of 1.5 mm/min. The exact depth of the slot was measured under a stereo microscope (Zeiss, Germany) using image analysis software (Soft-Imaging Software GmbH, Germany). The fracture toughness KIc was then calculated using the formula:

where Γm=1.9887−1.326a′−((3.49−0.68a′+1.35a′2)a′(1−a′))/(1+a′)2, Fmax = fracture load in N; L1, L2 = support distance; a′ = relative slot length a/W; B = specimen width in mm; W = specimen height in mm.

- 2.1.3. Chemical solubility

For the chemical solubility tests, 20 disc shaped specimens with a thickness of 1.2 ± 0.1 mm and a diameter of 15.0 ± 0.1 mm were used. The discs were cleaned in demineralized water and dried in an incubator for 4 h at 150 °C and then weighed on a precision scale to 0.1 mg accuracy. Then, the specimens were immersed in 4% acetic acid for 16 h at 80 °C. The cleaning and drying process was repeated and the specimens were weighed again. The initial and the final weights were used to calculate the chemical solubility using the formula:

where m0 = initial specimen weight; mn = weight after immersion in acetic acid; A = specimen surface (2πr2 + 2πrh).

- 2.2. Testing of CAD/CAM machined crowns

Sixty-four caries-free human molars were selected for similar size, cleaned by scaling and stored in 0.1% thymol solution throughout the course of the study. All teeth were visually examined and only those that were found to be sound and free from defects and cracks were used. Teeth were selected for the study if the variation in length and width was within 1 mm compared to the mean values (length 20 mm; diameter 10.5 mm). They were randomly assigned into 2 groups of 16 samples and 1 test group of 32 samples.

A representative clinical model with a missing first molar was selected [4] and [34]. The selected model was embedded using a silicone putty material (Formasil Xact, Heraeus Kulzer, Germany). The occlusal table of the model was set parallel to the horizontal plane. Afterwards, a silicone mold of this representative model was fabricated with a silicone impression material (Formasil Xact). This index was then used as the negative form for fixing all abutment teeth in the sample holder and as an aid for the wax-up of the final crowns.

It has been demonstrated that abutment mobility can be a decisive factor in the evaluation of fracture strength of ceramic materials [21]. In order to imitate physiological tooth mobility, all roots of the abutment teeth were covered with an artificial periodontal membrane made out of gum resin (Anti-Rutsch-Lack, Welenko, Germany) [22] and [25]. In order to conform to the biological width, each tooth was coronally covered with wax 2 mm short of the cemento-enamel junction and then dipped once in the gum resin. After the gum resin had dried, the excess resin on the root tip was removed using a scalpel so that a uniform coating remained on the root surface.

The natural teeth were fixed into the silicone mold, the sample holder was attached to the prepared silicone mold and self-curing polyester resin (Technovit 4000, Heraeus Kulzer, Germany) was mixed and poured into the sample holder. After the resin had set, the silicone mold was removed and the crowns of the abutment teeth were cleaned.

An occlusal reduction of 1.8–2.0 mm was prepared with diamond burs leaving an abutment height of approximately 4.5 mm, followed by a circular 1.2 mm wide shoulder. Teeth were prepared with a 6° convergence angle.

Light body silicon impression material (Monopren®, Kettenbach, Germany) was syringed around the abutments, and putty material (Monopren®) was used on the tray. After removing the impression from the abutment teeth, the master model was poured in type 4 dental stone (Everest Rock, KaVo, Germany). After a setting time of at least 45 min, the model was removed from the impression. Sectioned dies were prepared and two layers of dye spacer (Purargent) were applied.

- 2.2.1. ZrSiO4-crowns

Thirty-two test crowns were fabricated from Everest HPC ZrSiO4 ceramic. For this purpose, a full wax-up was fabricated for each crown using the silicone index as template. Then, the abutment teeth were scanned and a second scan was performed of the outer surface of the waxed crowns (KaVo Everest Scan, KaVo, Germany). The Everest CAD/CAM system was used to generate the data for the manufacturing of the crowns. Everest HPC blanks were mounted in the Everest Engine CAM milling unit and the crowns were milled in the green stage. Subsequently, a tempering process at a maximal temperature of 1575 °C and sintering time of up to 4 h was performed in the Everest therm unit to sinter the green bodies. During this reaction sintering process, first PMSS is pyrolyzed, then ZrSi2 is oxidized, and finally both ZrSiO4 is formed and the ceramic is sintered to dense bodies.

The finalized crowns were tried on the plaster dies and adjusted as required.

- 2.2.2. Porcelain-fused-to-metal crowns

The copings of all porcelain-fused-to-metal crowns crowns were waxed-up with a minimum thickness of 0.4–0.5 mm. After investing (GC-Vest) and burnout of the wax patterns all crowns were cast in a high precious gold alloy (V-Classic). After divesting, the frameworks were sandblasted and cleaned with jet steam, then heated in a furnace at 950 °C for 10 min to create an oxidation layer.

Vita Omega porcelain was used for veneering. The veneer thickness varied between 1.5 mm on the occlusal surface and 0.7–1 mm on the circumference.

- 2.2.3. IPS Empress 2 crowns

The copings of all Empress 2 control crowns (Ivoclar, Schaan, Switzerland) were fabricated with a wall thickness of at least 0.8 mm following the manufacturer's recommendations. The patterns were sprued and invested (IPS Empress investment). The automatic heat pressing process was performed in the EP 500 press furnace using standard parameters. After completion of the pressing and cooling cycle, rough divestment was carried out. Subsequently, the pressed copings were cleaned and the final tooth shape was built up using the IPS Empress 2 ceramic material and final staining and glazing were performed.

- 2.2.4. Luting procedure

Glass–ionomer cement (Ketac-Cem, 3M Espe, Germany) was used for the cementation of the ZrSiO4 and PFM restorations in this study. The inner surfaces of all crowns were air particle abraded with 100 μm aluminium oxide. The abutment teeth were cleaned with water spray for 30 s and dried with air for 60 s. Glass–ionomer cement (Ketac-Cem) was mixed and applied to the inner surfaces of all crowns using a small brush. The crowns were seated on the abutment teeth and held in place with finger pressure. Any excess cement was removed by wiping with foam pellets.

Variolink®II (Ivoclar) dual-curing resin cement was used for the cementation of the Empress 2 restorations in this study. The inner surfaces of Empress 2 crowns were etched for 20 s with IPS Empress Ceramic etch gel, rinsed with water for 30 s and dried. A silane coupling agent was applied with a brush and after 60 s dried with air. All abutments were then etched with 37% phosphoric acid for 30–60 s. They were cleaned with water spray and dried with air. Syntac Primer (Ivoclar) was applied and dispersed after 15–20 s with an air syringe. Syntac adhesive (Ivoclar) was applied and dispersed with air after 10 s. Heliobond (Ivoclar) was applied and dispersed with air. At the same time Variolink II was mixed and applied evenly on the inner surfaces of all retainers; then the crowns were seated on the abutments and held in place with finger pressure. All surfaces of the Empress 2 crowns were light-cured for 1 min.

- 2.2.5. Fatigue tests: thermomechanical loading

All specimens in each group were exposed to 1.2 million cycles of thermo-mechanical fatigue in a computer-controlled chewing simulator simulating clinically relevant conditions (Willytec, Germany) [26]. A load of 49 N was applied in the center of the occlusal surface of the crowns using 6 mm diameter ceramic antagonist ball (Steatit®, Ceramtec, Germany) [8] and [26]. The vertical movement of the antagonist was set at 6 mm and the horizontal way at 0.5 mm, which resulted in a cycle frequency of 1.3 Hz. During testing, all samples were allowed to reach a thermal equilibrium between 5 and 55 °C for 60 s each with an intermediate pause of 12 s, maintained by a thermostatically controlled liquid circulator.

During the dynamic loading, all samples were examined twice a day. Fractures of the tooth or any part of the restoration were recorded as a failure.

All specimens were thereafter loaded until fracture occurred using a universal testing machine (Zwick Z010/TN2S, Zwick, Germany). A 1 mm thick tin foil was placed over the occlusal surface to achieve a homogeneous stress distribution. A perpendicular load was applied to the occlusal surface of the molar at a crosshead speed of 2 mm/min. The loads required to fracture the samples were recorded with the Zwick testXpert® V 7.1 software.

The statistical analysis of the fracture load tests was performed using the Kruskal–Wallis test for the comparison of multiple groups and pair-wise comparisons with the Wilcoxon test (SPSS, Germany) at a significance level 0.05.

- 3. Results

- 3.1. Material tests

- 3.1.1. Flexural strength

Flexural strength was measured using a biaxial bend setup. The mean diameter of the disk specimens (n = 12) was 15.1 mm and the average thickness was 1.2 mm. Based on the average fracture load of 281.7 N, a flexural strength of 328.3 MPa was calculated. Using the described approach, a Weibull strength of 338.6 MPa and a Weibull modulus m of 16.8 were found (Table 1).

Table 1.

Flexural strength and Weibull parameters of the novel ZrSiO4 ceramic

| n | 10 |
| --- | --- |
| Flexural strength, σ [MPa] | 328.3 |
| S.D. [MPa] | 22.5 |
| Weibull strength, σ0 [MPa] | 338.6 |
| Weibull modulus, m | 16.8 |

- 3.1.2. Fracture toughness

Fracture toughness was measured using box shaped specimens. The mean height and width of the specimens (n = 10) were 2.49 and 1.38 mm, respectively. The mean depth of the sawed slots was 1.31 mm (S.D. 0.01 mm). The average fracture load of the specimens was 19.38 N (S.D. 2.34 N) which resulted in a mean fracture toughness of 5.16 MPa m1/2 (S.D. 0.48 MPa m1/2; see Table 2).

Table 2.

Comparison of the flexural strengths of selected dental ceramic materials [13] and [31]

|  | Flexural strength, σ [MPa] | Weibull strength, σ0 [MPa] | Weibull modulus, m | Fracture toughness KIc [MPa m1/2] |
| --- | --- | --- | --- | --- |
| Zirconia-TZP | 913 | 937 | 18 | 9.42 |
| Everest HPC | 328 | 339 | 17 | 5.16 |
| In-Ceram alumina | 666 | 290 | 5 | 5.00 |
| IPS Empress 2 | 442 | 289 | 9 | 2.48 |
| Cerec Mark II | 86 | 88 | 24 | 1.29 |
| IPS Empress 1 | 84 | 89 | 9 | 1.17 |

- 3.1.3. Chemical solubility

For the chemical solubility tests, the 20 disk specimens were submitted to the tests as described earlier. The average weight loss for a single specimen was 291.5 μg. This resulted in a chemical solubility of 7.2 μg/cm2.

- 3.2. Testing of CAD/CAM machined crowns

All the test samples (n = 64) survived after 1,200,000 cycles of dynamic loading and thermocycling in the artificial oral environment. The mean fracture strengths (±S.D.) for the crowns were: 1789.6 N (658.9) for ZrSiO4, 1715 N (854.3) for Empress 2 crowns and 2415.5 N (801.8) for PFM crowns. The distribution of the values can be obtained from Fig. 1. The statistical comparison with the Kruskal–Wallis–test showed an overall significant difference between the fracture loads (P = 0.020). The multiple pair-wise comparisons showed significant differences between the fracture loads of PFM crowns and Empress 2 crowns (P = 0.032) as well as PFM and ZrSiO4-crowns (P = 0.007). However, there was no significant difference between ZrSiO4-crowns and Empress 2 crowns (P = 0.743; Mann–Whitney U-tests).

|  | Display Full Size version of this image (12K) |
| --- | --- |

Fig. 1. Boxplot of the fracture load of the three types of crowns after thermomechanical loading in the dual axis chewing simulator.

- 4. Discussion

In these manuscript data from in vitro tests of a new shrinkage-free all-ceramic material are reported.

A number of static material tests were performed to confirm the mechanical properties of this material under standardized conditions. A DIN EN ISO 6872 bend test was performed first. A flexural strength of 328 MPa was demonstrated in this study. While values are far lower than for dense sintered aluminium oxide ceramics or zirconium oxide, this value puts the ZrSiO4 ceramic Evererst HPC in close proximity to other ceramics which involve a sintering process, such as IPS Empress 2. For a comparison of the flexural strengths of different ceramic materials used for dental restorations, see Table 2. This strength makes the material suitable for single tooth restorations of molars and premolars. The high Weibull strength (339 MPa, see Table 2) and the high Weibull modulus [17] further emphasize the suitability of this material for use in single crown applications [13] and [31].

The obtained fracture toughness of the tested ZrSiO4 ceramic was relatively high. The value of slightly above 5 MPa m1/2 was comparable to In-Ceram alumina (see Table 2), another oxide based ceramic material [13] and [31]. While we found a low standard deviation indicating stable results and an adequate sample size there still may be limitations to our data. We had a relatively high notch width which may result a large notch radius. The measured fracture toughness is dependent on the square root of the notch radius [33] thus higher radiuses may result in higher fracture toughness values. The fracture toughness values as found in this study should be interpreted with caution.

The tested ZrSiO4 ceramic also possesses a low chemical solubility. The solubility of 7.2 μg/cm2 is well below a threshold of 10 μg/cm2 and thus it passes the DIN EN ISO 6872 benchmark of 100 μg/cm2 without problems. The chemical solubility of this ceramic is comparable to aluminium oxide and yttrium-stabilized zirconium oxide. A low chemical solubility is also imperative for a good biocompatibility. The absence of cytotoxicity, hemotoxicity and genotoxicity have also been confirmed in independent tests following DIN EN ISO 10993-1 [12].

Fracture strength tests of ceramic materials are important for the expected lifetime with an acceptable low probability of failure [35]. The fitness of a new ceramic material for a range of indications can be decided upon by comparing results gained in in vitro studies with values for known and accepted materials. Therefore, the physical properties of newly developed dental materials must be tested before they can be recommended for clinical use. However, the characteristics of clinical failure of all-ceramic restorations are not well reflected by parameters such as flexural strength of material specimens or static load-to-failure tests of complete restorations. These tests do not well replicate the failure mechanisms seen in retrieved clinical specimens [20]. Therefore, it is more clinically relevant to test the specimens under physiologic fatigue load in a chewing simulator. This allows the evaluation of dental restorative systems under clinically relevant conditions while saving evaluation time [4], [9], [20], [23], [25], [26] and [43]. It has been shown that different all-ceramic restorations show different fatigue behaviors. The chewing machine used in this study was developed to reproduce the in vivo environment by adding moisture and a controlled temperature to the test conditions [23]. This is important when measuring fracture or fatigue strength of dental ceramics, as water and temperature change can act at crack tips to decrease the strength of glass and ceramics [20] and [35]. According to the literature, fatigue produced by 240,000 cycles in the chewing simulator corresponds to a period of 1 year of clinical service [37]. Therefore, 1,200,000 chewing cycles were performed to simulate a service time of 5 years [23] and [26]. The parameters used for the chewing machine were limited to the physiological values found in the literature. While maximum bite forces can easily achieve in excess of 500 N in the posterior areas, regular masticatory forces can, but do not generally exceed the 49 N loading force used in this study [8], [15] and [38].

Clinical studies are not indicated before the safety of a restorative material has been confirmed by in vitro tests. In addition, they are time consuming, expensive and the standardization of the test parameters is difficult. Therefore, more valid tests such as the chewing simulation carried out in this study are required to include other parameters that are incurred by restorations in the mouth.

All crowns survived the chewing simulation, resulting in a survival of 100% of the specimens. Thereafter, all crowns were loaded to failure in a universal testing machine. The fracture loads obtained by this test only reflect the loads at which failure of the crowns occurs. These values varied between 966 and 3964 N. These loads should not be confused with the values obtained by the static tests on material specimens. Crown failure loads strongly depend on the thickness and geometry of the crowns.

The fracture load of the Empress 2 control group shows that crowns made from Everest HPC exhibit comparable fracture loads; however, the loads for Everest HPC were obtained with conventional cementation. The high fracture loads obtained after cyclic loading thus emphasize that conventional cementation – as performed for these tests – appears as a viable option. Published data on bite forces indicate that the maximum biting force that may occur in the posterior dental area varies between 300 and 880 N [2] and [15]. In the present study all specimens achieved fracture strength values above 950 N. These values exceed the minimum fracture load limit (500 N) suggested by Kappert et al. [18] and [19].

Clinical survival rates between 95 and 97% over a period of up to 7 years can be expected of conventional (cast and PFM) fixed restorations [6], [24] and [27] and after 15–20 years survivals rates of 75% can be expected [7] and [39]. These survival rates and observation times so far well exceed those of any all-ceramic system. Clinical success of Empress 2 and In-Ceram all-ceramic restorations are promising at survival rates of 78–98% over 3–6 years [30], [32] and [40]; however, true long-term data or meta analyses as for conventional fixed restorations are absent.

In this study, extracted human teeth were used as abutments because their modulus of elasticity, bonding characteristics, thermal conductivity and strength are closer to the clinical situation than those of metal, plastic and animal teeth. Furthermore, natural teeth allowed the adhesive cementation of the specimens. An attempt was made to select abutments with comparable sizes. The extracted human teeth were stored in 0.1% thymol solution, preventing them from drying out and thereby becoming brittle [16] and also inhibiting microbial activity in the storage solution [42].

The following conclusions were drawn:

1. The ZrSiO4 ceramic tested in this study possesses mechanical properties that are comparable to other all-ceramic dental materials.

2. Full crowns made from the zircon ceramic withstand masticatory forces found in the posterior area.

3. While the material appears to be reliable for clinical use for posterior full crowns, clinical assessment of these all-ceramic restorations is required.

- References

[1] M.A. Awad and J.S. Feine, Measuring patient satisfaction with mandibular prostheses, Community Dent Oral Epidemiol 26 (1998), pp. 400–405. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (35)

[2] J.F. Bates, G.D. Stafford and A. Harrison, Masticatory function—a review of the literature. 2. Speed of movement of the mandible, rate of chewing and forces developed in chewing, J Oral Rehabil 2 (1975), pp. 281–301. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (29)

[3] A.J. Beaumont Jr., An overview of esthetics with removable partial dentures, Quintessence Int 33 (2002), pp. 747–755. View Record in Scopus | Cited By in Scopus (2)

[4] M. Behr, M. Rosentritt, A. Leibrock, S. Schneider-Feyrer and G. Handel, In vitro study of fracture strength and marginal adaptation of fibre-reinforced adhesive fixed partial inlay dentures, J Dent 27 (1999), pp. 163–168. SummaryPlus | Full Text + Links | PDF (830 K) | View Record in Scopus | Cited By in Scopus (37)

[5] J.R. Binder and W. Die Qual der, Materialvielfalt bei vollkeramischem Zahnersatz, ZWL 3 (2004), pp. 46–51.

[6] J. Coornaert, P. Adriaens and J De Boever, Long-term clinical study of porcelain-fused-to-gold restorations, J Prosthet Dent 51 (1984), pp. 338–342. Abstract | View Record in Scopus | Cited By in Scopus (38)

[7] N.H. Creugers, A.F. Kayser and M.A. van’t Hof, A meta-analysis of durability data on conventional fixed bridges, Community Dent Oral Epidemiol 22 (1994), pp. 448–452. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (81)

[8] J.A. De Boever, W.D. McCall Jr., S. Holden and M.M. Ash Jr., Functional occlusal forces: an investigation by telemetry, J Prosthet Dent 40 (1978), pp. 326–333. Abstract | View Record in Scopus | Cited By in Scopus (54)

[9] R. DeLong and W.H. Douglas, Development of an artificial oral environment for the testing of dental restoratives: bi-axial force and movement control, J Dent Res 62 (1983), pp. 32–36. View Record in Scopus | Cited By in Scopus (93)

[10] W.J. Dunn, D.F. Murchison and J.C. Broome, Esthetics: patients’ perceptions of dental attractiveness, J Prosthodont 5 (1996), pp. 166–171. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (22)

[11] R.L. Engelmeier, Complete-denture esthetics, Dent Clin North Am 40 (1996), pp. 71–84. View Record in Scopus | Cited By in Scopus (7)

[12] Fiedler S, Ohmacht M. Zertifikat Biokompatibilität des Prüfmaterials KaVo Everest HPC. Ochsenhausen: MDT Medical Device Testing; 2005.

[13] H. Fischer and R. Marx, Festigkeit von Dentalkeramik, ZWR 109 (2000), pp. 240–244.

[14] R. Gerhard and N. Sawyer, Dentures to harmonize with heavily pigmented tissues, J Am Dent Assoc 73 (1966), pp. 94–95. View Record in Scopus | Cited By in Scopus (3)

[15] C.H. Gibbs, P.E. Mahan, A. Mauderli, H.C. Lundeen and E.K. Walsh, Limits of human bite strength, J Prosthet Dent 56 (1986), pp. 226–229. Abstract | View Record in Scopus | Cited By in Scopus (47)

[16] A.R. Helfer, S. Melnick and H. Schilder, Determination of the moisture content of vital and pulpless teeth, Oral Surg Oral Med Oral Pathol 34 (1972), pp. 661–670. Abstract | View Record in Scopus | Cited By in Scopus (51)

[17] V.D. Hennige, J. Hausselt, H.J. Ritzhaupt-Kleissl and T. Windmann, Shrinkage-free ZrSiO4-ceramics: characterisation and applications, J Eur Ceram Soc 19 (1999), pp. 2901–2908. SummaryPlus | Full Text + Links | PDF (270 K) | View Record in Scopus | Cited By in Scopus (9)

[18] H.F. Kappert, H. Knode and L. Manzotti, Metal-free bridges in posterior region, Dent Labor (Munch)38 (1990), pp. 177–183. View Record in Scopus | Cited By in Scopus (8)

[19] H.F. Kappert, H. Knode and R. Schultheiss, Strength of the In-Ceram system under mechanical loads and thermocycling in artificial saliva, Dtsch Zahnarztl Z 46 (1991), pp. 129–131. View Record in Scopus | Cited By in Scopus (21)

[20] J.R. Kelly, Clinically relevant approach to failure testing of all-ceramic restorations, J Prosthet Dent 81 (1999), pp. 652–661. SummaryPlus | Full Text + Links | PDF (296 K) | View Record in Scopus | Cited By in Scopus (116)

[21] J.R. Kelly, J.A. Tesk and J.A. Sorensen, Failure of all-ceramic fixed partial dentures in vitro and in vivo: analysis and modeling, J Dent Res 74 (1995), pp. 1253–1258. View Record in Scopus | Cited By in Scopus (93)

[22] M. Kern, W.H. Douglas, T. Fechtig, J.R. Strub and R. DeLong, Fracture strength of all-porcelain, resin-bonded bridges after testing in an artificial oral environment, J Dent 21 (1993), pp. 117–121. Abstract | View Record in Scopus | Cited By in Scopus (26)

[23] M. Kern, J.R. Strub and X.Y. Lu, Wear of composite resin veneering materials in a dual-axis chewing simulator, J Oral Rehabil 26 (1999), pp. 372–378. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (49)

[24] T. Kerschbaum and R. Voss, Comparison of cast and metal–ceramic veneer crowns: results of a follow-up study in wearers of partial prostheses, Dtsch Zahnarztl Z 32 (1977), pp. 200–206. View Record in Scopus | Cited By in Scopus (4)

[25] S.O. Koutayas, M. Kern, F. Ferraresso and J.R. Stub, Influence of design and mode of loading on the fracture strength of all- ceramic resin-bonded fixed partial dentures: an in vitro study in a dual-axis chewing simulator, J Prosthet Dent 83 (2000), pp. 540–547. SummaryPlus | Full Text + Links | PDF (265 K) | View Record in Scopus | Cited By in Scopus (15)

[26] I. Krejci, T. Reich, F. Lutz and M. Albertoni, An in vitro test procedure for evaluating dental restoration systems. 1. A computer-controlled mastication simulator, Schweiz Monatsschr Zahnmed 100 (1990), pp. 953–960. View Record in Scopus | Cited By in Scopus (42)

[27] P.J. Leempoel, S. Eschen, A.F. De Haan and M.A. Van’t Hof, An evaluation of crowns and bridges in a general dental practice, J Oral Rehabil 12 (1985), pp. 515–528. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (42)

[28] R.G. Luthardt, M. Holzhuter and O. Sandkuhl et al., Reliability and properties of ground Y-TZP-zirconia ceramics, J Dent Res 81 (2002), pp. 487–491. View Record in Scopus | Cited By in Scopus (24)

[29] R.G. Luthardt, M.S. Holzhuter, H. Rudolph, V. Herold and M.H. Walter, CAD/CAM-machining effects on Y-TZP zirconia, Dent Mater 20 (2004), pp. 655–662. SummaryPlus | Full Text + Links | PDF (433 K) | View Record in Scopus | Cited By in Scopus (16)

[30] P. Marquardt and J.R. Strub, Survival rates of IPS Empress 2 all-ceramic crowns and fixed partial dentures: results of a 5-year prospective clinical study, Quintessence Int 37 (2006), pp. 253–259. View Record in Scopus | Cited By in Scopus (11)

[31] R. Marx, H. Fischer, M. Weber and F. Jungwirth, Crack parameters and Weibull moduli: subcritical crack growth and long-time durability of all-ceramic materials, Dtsch Zahnärztl Z 56 (2001), pp. 90–98.

[32] E.A. McLaren and S.N. White, Survival of In-Ceram crowns in a private practice: a prospective clinical trial, J Prosthet Dent 83 (2000), pp. 216–222. Abstract | PDF (692 K) | View Record in Scopus | Cited By in Scopus (40)

[33] D. Munz and T. Fett, Fracture mechanics. In: D. Munz and T. Fett, Editors, Ceramics: mechanical properties, failure behaviour, materials selection, Springer, Berlin, Heidelberg (1999), pp. 19–51.

[34] W. Riedling and H.F. Kappert, Longevity of Hi-Ceram full ceramic bridges, Quintessenz 38 (1987), pp. 1493–1498. View Record in Scopus | Cited By in Scopus (1)

[35] J.E. Ritter, Predicting lifetimes of materials and material structures, Dent Mater 11 (1995), pp. 142–146. SummaryPlus | Full Text + Links | PDF (617 K) | View Record in Scopus | Cited By in Scopus (25)

[36] A.S. Rizkalla and D.W. Jones, Mechanical properties of commercial high strength ceramic core materials, Dent Mater 20 (2004), pp. 207–212. SummaryPlus | Full Text + Links | PDF (338 K) | View Record in Scopus | Cited By in Scopus (10)

[37] R.L. Sakaguchi, W.H. Douglas, R. DeLong and M.R. Pintado, The wear of a posterior composite in an artificial mouth: a clinical correlation, Dent Mater 2 (1986), pp. 235–240. Abstract | View Record in Scopus | Cited By in Scopus (44)

[38] H. Schwickerath and I. Coca, Single crowns of glass–ceramic, Phillip J Restaur Zahnmed 4 (1987), pp. 336–338. View Record in Scopus | Cited By in Scopus (4)

[39] M.S. Scurria, J.D. Bader and D.A. Shugars, Meta-analysis of fixed partial denture survival: prostheses and abutments, J Prosthet Dent 79 (1998), pp. 459–464. SummaryPlus | Full Text + Links | PDF (158 K) | View Record in Scopus | Cited By in Scopus (74)

[40] B.S. Segal, Retrospective assessment of 546 all-ceramic anterior and posterior crowns in a general practice, J Prosthet Dent 85 (2001), pp. 544–550. Abstract | PDF (193 K) | View Record in Scopus | Cited By in Scopus (24)

[41] P.N. Sellen, D.C. Jagger and A. Harrison, Methods used to select artificial anterior teeth for the edentulous patient: a historical overview, Int J Prosthodont 12 (1999), pp. 51–58. View Record in Scopus | Cited By in Scopus (14)

[42] O. Sparrius and E.S. Grossman, Marginal leakage of composite resin restorations in combination with dentinal and enamel bonding agents, J Prosthet Dent 61 (1989), pp. 678–684. Abstract | View Record in Scopus | Cited By in Scopus (10)

[43] J.R. Strub and S.M. Beschnidt, Fracture strength of 5 different all-ceramic crown systems, Int J Prosthodont 11 (1998), pp. 602–609. View Record in Scopus | Cited By in Scopus (29)

Corresponding author at: Department of Prosthodontics, University Hospital, Hugstetter Sr. 55, D-79106 Freiburg, Germany. Tel.: +49 761 270 4906; fax: +49 761 270 4824.

Dental Materials Volume 23, Issue 7, July 2007, Pages 785-791

Material características de um romance encolhimento - livre ZrSiO4 cerâmica para a fabricação de coroas póstero

G. Heydeckea, b,,, F. Butza, J. R. Binderc e J. R. Struba 
ADepartment de Prostodontia, University Hospital, Freiburg, Alemanha 
BFaculty de Odontologia, McGill University, Montreal, Canadá 
CInstitute para Materials Research III, Forschungszentrum Karlsruhe, Karlsruhe, Alemanha 
Recebeu 15 de Agosto de 2005; Revista 21 March 2006; Aceite 22 de Junho de 2006. Disponível online 1 de Setembro de 2006.

Abstract 
Objectivo 
Para confirmar a aplicabilidade clínica de um romance ZrSiO4 (zircon), com base encolhimento - livre caco, a flexurais força, fratura dureza e solubilidade química foram testados. Além disso, a fractura de plena carga coroas feitas a partir deste material foi testado após cíclico thermomechanical carregamento.

Métodos 
Flexurais força de 12 espécimes foi medida usando um biaxial vincos teste. Fracture dureza foi medido utilizando 10 fendas caixa moldada espécimes. Os espécimes foram fraturados usando uma máquina ensaios universais; Fratura cargas foram registrados. A química solubilidade teste foi realizado em conformidade com a norma ISO norma 6872. Além disso, 32 ZrSiO4 todos - cerâmicos coroas foram fabricados em extraído caries - molares humanos livres. Dezesseis Empress 2 e 16 PFM coroas serviu como uma referência controle. Após artificiais envelhecimento de 1,2 milhões de ciclos no roer simulador, a taxa de sobrevivência das coroas foi determinada. A fratura carga de todos os exemplares sobreviventes foram obtidas mediante o carregamento coroas até fratura em uma máquina testes universais.

Resultados 
A flexurais força de 328 MPa foi encontrado. A fratura dureza da ZrSiO4 cerâmica foi 5,16 MPa m0.5. A química solubilidade elevou - se a 7,2 µ µ g / centímetro 2. Todos os espécimes sobreviveu a roer simulação (taxa de sobrevivência: 100%); Coroas não tinha que ser re - cimento. A média fratura força de 1790 N foi encontrado no Everest HPC para Empress 2 coroas, 1715 N para Empress 2 coroas e 2416 N para PFM coroas. Fracture cargas de PFM coroas foram significativamente mais elevados do que para Empress 2 coroas (P = 0,032), bem como ZrSiO4 - coroas (P = 0,007). Não houve diferença significativa entre ZrSiO4 - coroas e Empress 2 coroas (P = 0,743).

Significância 
Na fase actual, Everest HPC pode ser recomendado para a fabricação de coroas único como uma alternativa ao convencional PFM e outros - todos cerâmica coroas, devido a sua resistência fratura excede média masticatory forças na região posteriores. 


Palavras - chave: All - cerâmica; Fracture força; Coroas; Fadiga carregamento 

Artigo Outline 
1. Introdução 
2. Materiais e métodos 
2,1. Material testes 
2.1.1. Flexurais força 
2.1.2. Fracture dureza 
2.1.3. Chemical solubilidade 
2,2. Teste de CAD / CAM maquinadas coroas 
2.2.1. ZrSiO4 - coroas 
2.2.2. Porcelana - fundiram - a - metálicos coroas 
2.2.3. IPS Empress 2 coroas 
2.2.4. Reboco procedimento 
2.2.5. Fadiga testes: thermomechanical carregamento 
3. Resultados 
3,1. Material testes 
3.1.1. Flexurais força 
3.1.2. Fracture dureza 
3.1.3. Chemical solubilidade 
3,2. Teste de CAD / CAM maquinadas coroas 
4. Discussão 
Referências 

1. Introdução 
Mastigatórios eficiência, fonética e aparência estética são essenciais determinantes do doente satisfação com próteses dentárias [1]. No entanto, o indivíduo avaliação de estética é muito subjetiva [41]. Ao longo das últimas décadas, as expectativas do paciente dentária estética tem vindo a aumentar continuamente, particularmente no que se relacionam com a selecção de dente cor [10]. Um elevado grau de sofisticação foi alcançado no reino do dente cor estética. Três fatores podem ser responsabilizados por esse desenvolvimento: (1) a influência de técnicos altamente qualificados dentária, (2) o contínuo aperfeiçoamento das técnicas fabrico, e (3) a melhoria dos materiais dentários [3], [11] e [14] . No entanto, estéticas alternativas à plena copa ouro, como porcelana - fundiram - a - metálicos (PFM) coroas são geralmente mais caros. All - cerâmicos coroas são ainda mais intensiva dos recursos no que respeita à sua produção e custo. Sistemas de segurança social abrange principalmente as menos onerosas opção tornando - o elenco ouro coroa o padrão de cuidados. 
Para além do PFM técnica, há uma série de diferentes sistemas disponíveis para o fabrico de todas as restaurações de cerâmica. No entanto, todos esses sistemas foram introduzidas com o objectivo de melhorar a aparência do dente coloridas fixas prosthodontics e alargar a sua indicação para a parte posterior área. Estes podem ser considerados topo final restaurações. Há poucos esforços no sentido de melhorar a estética em nível normal. CAD / CAM técnicas parecem estar a mais adequada para a criação de custo efetivo fixo restaurações. Mas, até agora, não foram exploradas com todo o seu potencial para fornecer simples e de baixo custo único coroas. Se perguntar doentes, a maioria irá responder que são dentes naturais "branco". Embora isso possa parecer um grosseiros definição de um dentista profissional este poderia ser um ponto de partida para alternativas simples para um elenco ouro coroa, que não é branco. 
Várias técnicas têm sido descritas para a fabricação de todos - cerâmico fixo restaurações. Por exemplo, com óxido de alumínio vidro infiltradas núcleo cerâmicas (In - Ceram Alumina, Vita, Alemanha) uma força de flexurais> 450 MPa pode ser alcançado. Com calor pressionado lítio - Disilicato núcleos, flexurais fortes do ligeiramente inferior a 400 MPa, foram relatados [36]. Ambos os tipos de materiais necessários altamente qualificados manuais para os respectivos cueca - vazamento, cera ou calor pressionando procedimentos. Uma das desvantagens de todas as suas sinterizada cerâmica é linear de sinterização retração de até 20%. Para obter o desejado dimensões, restaurações deve ser formada para correspondentemente alargada dimensões. 
Mais recente fabricação técnicas envolvem a usinagem de CAD - concebidas restaurações (e.g., Cerec, Sirona, Alemanha; DCS, Suíça; Cercon; Degudent, Alemanha). O material é mais prevalente núcleo Y - TZP zircónia. A usinagem de restaurações de sinterizados ZrO2 é bastante morosa; O processo de usinagem uma única unidade restauração freqüentemente leva 2-3 horas [5]. Além disso, a extrema dureza do sinterizada alta resistência cerâmica resultados em alto desgaste das ferramentas usinagem. Dados recentes revelam que a usinagem processo também tem um efeito negativo sobre a microestrutura do material. Como consequência, a induzida microcracks reduzir a fiabilidade das restaurações de usinagem totalmente sinterizada ZrO2 [28] e [29]. Restaurações feitas utilizando estas técnicas requerem uma veneering camada principalmente em (a) conseguir um bom dente forma e (b) para cobrir os materiais essenciais para alcançar aceitável estética. 
Uma alternativa para cueca - vazamento ou trituração de blocos sinterizados é CAD / CAM usinagem de pré - sinterizada blocos ou em branco no estádio verde. A utilização de um convencional caco processo exige um corpo alargada verde para compensar a retração linear de sinterização. O alargamento fator depende da densidade do compacto verde, a densidade do corpo após sinterização e seus respectivos pesos. Assim, a precisão da cor branca directamente depende de uma qualidade constante dos órgãos verde ea estabilidade da sinterização processo. 
Estas áreas problema pode ser contornado se um encolhimento - livre caco é usado. Com uma chamada reação vinculação processo, isto é conseguido através da expansão do volume de um componente durante o processo sinterização compensar o volume perda de outro. Se uma baixa perda aglutinante é adicionado, a sinterização retração pode ser reduzido ainda mais. Tal sistema baseado em zircónio silicide (ZrSi2), zircónia (ZrO2) e um polymethylsilsesquioxane (PMSS), foi descrito em detalhe noutros países [17]. A cerâmica baseada neste sistema é zircon (ZrSiO4). Ela pode ser moldado para espaços em branco utilizando axial ou grafites isostáticas pressionando. Os espaços em branco são estáveis e podem ser trabalhadas utilizando comercial CAD / CAM equipamentos. 
Os objetivos deste estudo in vitro foram as seguintes: 
1. - Se caracterizar a zircon (ZrSiO4) mais testes com procedimentos pertinentes ao dentista cerâmica; 
2. Para avaliar a taxa de sobrevivência de todas as coroas de cerâmica fabricados a partir zircon cerâmica, após 1,2 milhões de ciclos de carga em um duplo eixo roer simulador; 
3. Para medir a carga de fratura de todas as coroas de cerâmica fabricados a partir zircon cerâmica, após 1,2 milhões de ciclos de carga em um duplo eixo roer simulador. 
2. Materiais e métodos 
Para testar o novo ZrSiO4 cerâmica (HPC Everest, KaVo, Biberach, Alemanha), no que respeita à sua adequação para restaurações dentárias, uma série de testes foram realizados. Isso incluiu flexurais força ensaio, o cálculo do Weibull módulo, a medição da fratura dureza, assim como a química solubilidade. 
2,1. Material testes 
2.1.1. Flexurais força 
Flexurais força foi medida usando um biaxial pistão - em - 3 - bola teste. Doze disco moldado espécimes foram preparados com um diâmetro de 15,1 ± 0,1 mm e uma espessura de 1,2 ± 0,1 milímetros. Os espécimes foram medidos com um parafuso micrómetros e suas medidas reais registados. Em seguida, os espécimes foram colocados no dispositivo testes e ensaios, em seguida, foi realizado em uma máquina universais ensaios (Zwick, Alemanha), numa crosshead velocidade de 1,5 milímetros / minuto. O flexurais força σ foi calculada pela fórmula: 

  

Onde; Υ = 0,19 (Poisson's parâmetro de cerâmica); R1 = 5,97 milímetros (raio do círculo das bolas de apoio); R2 = 0,76 milímetros (raio do pistão); R3 = raio de O espécime em mm; Fmax = fratura carga em N; D = espessura do espécime em mm. 
Além disso, o Weibull força σ0 e os Weibull modulo milhões foram calculados utilizando análise de regressão linear. A fratura probabilidade P σ é definida pela seguinte equação: P = σ (n. - 0,5) / N, onde N é o número total de amostras e n refere - se à ordem crescente da σ - valores. Os logaritmos naturais são calculados de σ e entrou como x - valores. O correspondente y - valores são calculados utilizando a fórmula ln (ln (1 / (1 - P σ))) e, em seguida, são inseridos em coordenar o plano. O diagrama mostra uma linha reta com o declive m e a intersecção com os x - eixo ln σ0. 
2.1.2. Fracture dureza 
Fracture dureza foi medido utilizando 10 caixa moldada espécimes medindo 20 milímetros × 3 milímetros × 1,5 milímetros. Antes do ensaio, todas as medidas dos espécimes foram registrados. Em seguida, um entalhe foi preparada dentro de cada espécime até uma profundidade de 50% utilizando um diamante jig viu. A largura do entalhe foi definida pela espessura da lâmina viu cerca de 300 µ m de Os espécimes foram, em seguida, fraturado no universais testes máquina em uma crosshead velocidade de 1,5 milímetros / minuto. A exata profundidade da ranhura foi medido sob um microscópio estéreo (Zeiss, Alemanha), utilizando softwares gráficos análise (Soft - Software Imaging GmbH, Alemanha). A fratura dureza KIc foi então calculada pela fórmula: 

  

Onde Γ metros = 1,9887 - 1.326a '- ((3,49 - 0.68a' +1,35 a'2) um '(1 - um'))/( 1 + um') 2, Fmax = fratura Carga em N; L1, L2 = apoio distância; Uma '= parente ranhura comprimento de um / W; B = espécime largura em milímetros; W = espécime altura em mm. 
2.1.3. Chemical solubilidade 
Para os ensaios químicos solubilidade, 20 disco moldado espécimes com uma espessura de 1,2 ± 0,1 mm e um diâmetro de 15,0 ± 0,1 milímetros foram utilizados. Os discos foram limpos em água desmineralizada e secagem em uma incubadora de 4 horas a 150 ° C e, em seguida, pesava sobre a precisão escala de 0,1 mg de precisão. Em seguida, os espécimes foram imersos em 4% de ácido acético por 16 horas a 80 ° C. A limpeza e secagem processo foi repetido e os espécimes foram pesadas novamente. A primeira ea última pesos foram utilizados para calcular o produto químico solubilidade usando a fórmula: 

  

Onde m0 = massa inicial espécime; Sm = peso após imersão em ácido acético; A = espécime superfície (2 π r2 + 2 π rh). 
2,2. Teste de CAD / CAM maquinadas coroas 
Sessenta e quatro caries - livre molares humanos foram selecionados para a dimensão semelhante, limpos por escamação e armazenados em 0,1% timol solução durante todo o curso do estudo. Todos os dentes foram examinados visualmente e só aqueles que foram encontrados para ser sadia e livre de defeitos e rachaduras foram utilizados. Dentes foram selecionados para o estudo, se a variação no comprimento e largura foi dentro de 1 milímetro, em comparação com os valores médios (comprimento 20 mm, Diâmetro 10,5 milímetros). Eles foram aleatoriamente atribuídos em 2 grupos de 16 amostras e 1 teste grupo de 32 amostras. 
Um representante clínica modelo com uma falta primeiro molares foi selecionado [4] e [34]. O modelo selecionado foi incorporado utilizando um Silicone putty materiais (Formasil Xact, Heraeus Kulzer, Alemanha). A mesa oclusal do modelo foi fixado paralelo ao plano horizontal. Em seguida, um representante Silicone molde deste modelo foi fabricado com uma impressão Silicone materiais (Formasil Xact). Este índice foi então usado como negativa a forma de fixação todos abutment dentes na amostra titular e como um auxílio para a cera - até finais do coroas. 
Foi demonstrado que abutment mobilidade pode ser um factor decisivo na avaliação da fratura força de materiais cerâmicos [21]. A fim de imitar fisiológicas dente mobilidade, todas as raízes do abutment dentes foram cobertos com uma membrana periodonto artificiais feitas de goma resina (Anti - Rutsch - Lack, Welenko, Alemanha) [22] e [25]. A fim de obedecer ao biológicos largura, cada dente foi coronally cobertos com cera 2 milímetros aquém da junção cemento - esmalte e, em seguida, uma vez mergulhou na goma resina. Após a goma resina tinha secado, o excesso de resina sobre a raiz ponta foi removida utilizando um escalpelo, a fim de que um modelo revestimento permaneceu na raiz superfície. 
Os dentes naturais foram fixadas em Silicone o molde, o titular amostra foi anexada à preparado Silicone molde e auto - cura de resina de poliéster (Technovit 4000, Heraeus Kulzer, Alemanha) foi misturado e vertido na amostra titular. Após a resina tinham estabelecido, o Silicone molde foi removido e as coroas do abutment dentes foram limpos. 
Oclusal uma redução de 1.8-2.0 milímetros foi preparado com diamante burs deixando uma abutment altura de cerca de 4,5 milímetros, seguido por uma circular 1,2 mm de largura ombro. Dentes foram preparados com um 6 ° convergência ângulo. 
Light corpo silício impressão materiais (Monopren ®, Kettenbach, Alemanha) foi syringed redor do abutments, e putty materiais (Monopren ®) foi usado na bandeja. Após a remoção da impressão do abutment dentes, o comandante modelo foi derramada no tipo 4 dentária pedra (Everest Rock, KaVo, Alemanha). Depois de um tempo de fixação pelo menos 45 minutos, o modelo foi removido da impressão. Seccionada morre foram preparados e duas camadas de tinta spacer (Purargent) foram aplicados. 
2.2.1. ZrSiO4 - coroas 
Trinta e duas coroas teste foram fabricados a partir Everest HPC ZrSiO4 cerâmica. Para este efeito, a plena cera - até foi fabricado para cada coroa utilizando o índice Silicone como modelo. Então, o abutment dentes foram digitalizados e uma segunda varredura foi feita da superfície exterior do waxed coroas (KaVo Everest Scan, KaVo, Alemanha). O Everest CAD / CAM sistema foi usado para gerar os dados para o fabrico das coroas. Everest HPC espaços foram montados no Everest Engine CAM moagem unidade e as coroas foram milled no estádio verde. Posteriormente, um processo a uma têmpera máxima temperatura de 1575 º C e tempo de sinterização até 4 horas foi realizada no Everest therm unidade de sinterização o verde órgãos. Durante esta reação sinterização processo, em primeiro lugar PMSS é pyrolyzed, então ZrSi2 é oxidado e, finalmente, tanto ZrSiO4 é formado e as cerâmica é sinterizada a densa órgãos. 
O finalizado coroas foram julgados no gesso morre e ajustada conforme necessário. 
2.2.2. Porcelana - fundiram - a - metálicos coroas 
Os copings de todos porcelana - fundiram - a - metálicos coroas coroas foram waxed - se com uma espessura mínima de 0.4-0.5 milímetros. Depois de investir (GC - Vest) e queimados da cera padrões todas coroas foram expressos em um alto precioso ouro liga (V - Classic). Após divesting, os quadros foram sandblasted e limpa com injeção de vapor, então aquecida em um forno a 950 º C por 10 minutos para criar uma camada de oxidação. 
Vita Omega porcelana foi utilizado para veneering. O verniz espessura variou entre 1,5 milímetros no oclusal superfície e 0.7-1 milímetro na circunferência. 
2.2.3. IPS Empress 2 coroas 
Os copings de todos Empress 2 controle coroas (Ivoclar, Schaan, Suíça), foram fabricados com uma parede de espessura mínima de 0,8 milímetros seguir as recomendações do fabricante. Os padrões foram sprued e investido (IPS Empress investimento). O calor automático pressionando processo foi realizado no PE 500 prima forno utilizando parâmetros normais. Após a conclusão do ciclo pressionando e frio, áspero alienação foi efectuada. Posteriormente, o pressionado copings foram limpas e os dentes forma definitiva foi construída através do IPS Empress 2 caco e finais de coloração e vidros foram realizadas. 
2.2.4. Reboco procedimento 
Glass - ionomer cimento (Ketac - Cem, 3M Espe, Alemanha) foi utilizado para a cimentação do ZrSiO4 e PFM restaurações neste estudo. As superfícies interiores de todas as coroas foram ar partículas desgaste com 100 µ m de óxido de alumínio. O abutment dentes foram limpos com água pulverizada por 30 s e secos com ar por 60 seg Glass - ionomer cimento (Ketac - Cem) foi misturado e aplicado ao interior superfícies de todas as coroas usando um pequeno pincel. As coroas estavam sentados sobre o abutment dentes e mantidos em posição com o dedo pressão. Qualquer excesso de cimento foi removido por exterminando com espuma pastilhas. 
Variolink ® II (Ivoclar) dupla - cura de resina cimento foi utilizado para a cimentação do Empress 2 restaurações neste estudo. As superfícies interiores de Empress 2 coroas foram gravadas por 20 s com IPS Empress Ceramic sarge Gel, lavado com água para 30 s e secas. A silane acoplamento agente foi aplicado com uma escova e depois de 60 a seco com o ar. Todos abutments foram, em seguida, gravado com 37% de ácido fosfórico para 30-60 denomin Eles foram limpos e secos com água pulverizada com o ar. Syntac Primer (Ivoclar) foi aplicada e dispersa após 15-20 s com um ar seringa. Syntac adesivo (Ivoclar) foi aplicada e dispersa com ar após 10 seg Heliobond (Ivoclar) foi aplicada e dispersa com o ar. Ao mesmo tempo Variolink II foi misturado e aplicada uniformemente sobre as superfícies interiores de todos retainers; As coroas foram, em seguida, sentado sobre os abutments e mantidos em posição com o dedo pressão. Todas as superfícies do Empress 2 coroas luz - foram curados por 1 minutos. 
2.2.5. Fadiga testes: thermomechanical carregamento 
Todos os exemplares de cada grupo foram expostos a 1,2 milhões de ciclos de fadiga térmica - mecânico em um computador - controlado roer simulador simulando clinicamente relevantes condições (Willytec, Alemanha) [26]. A carga de 49 N foi aplicada no centro da superfície oclusal das coroas com 6 milímetros de diâmetro cerâmica antagonista bola (Steatit ®, Ceramtec, Alemanha) [8] e [26]. Os movimentos verticais do antagonista foi fixado em 6 mm e as horizontais forma a 0,5 milímetros, o que resultou em um ciclo frequência de 1,3 Hz. Durante o ensaio, todas as amostras foram autorizados a chegar a um equilíbrio térmico entre 5 e 55 ° C para 60 a cada um, com uma pausa de 12 a intermediária, mantido por um thermostatically controlada líquido circulator. 
Durante o carregamento dinâmico, todas as amostras foram analisadas duas vezes por dia. Fraturas do dente ou em qualquer parte do restauro foram registados como um fracasso. 
Todos os espécimes foram posteriormente carregado até fratura ocorreu usando uma máquina universais ensaios (Zwick Z010/TN2S, Zwick, Alemanha). A 1 mm de espessura folha de estanho foi colocado sobre a superfície oclusal para conseguir uma distribuição homogénea estresse. A carga foi aplicada perpendicularmente à superfície oclusal dos molares numa crosshead velocidade de 2 milímetros / minuto. As cargas necessárias para fratura das amostras foram gravadas com a Zwick testXpert ® V 7,1 softwares. 
A análise estatística da fratura carga testes foi realizada utilizando o teste Kruskal - Wallis para a comparação de vários grupos e par - sábio comparações com o teste de Wilcoxon (SPSS, Alemanha) com um significado nível 0,05. 
3. Resultados 
3,1. Material testes 
3.1.1. Flexurais força 
Flexurais força foi medida usando um biaxial dobre configuração. O diâmetro médio do disco espécimes (N = 12) foi de 15,1 mm ea espessura média foi 1,2 milímetros. Com base na média fratura carga de 281,7 N, uma flexurais força de 328,3 MPa foi calculada. Usando a abordagem descrita, uma Weibull força de 338,6 MPa e um Weibull modulo m de 16,8 foram encontrados (Tabela 1). 
Tabela 1. 
Flexurais força e Weibull parâmetros do romance ZrSiO4 cerâmica 
N. 10 
Flexurais força, σ [MPa] 328,3 
S.D. [MPa] 22,5 
Weibull força, σ0 [MPa] 338,6 
Weibull módulo, metros 16,8 
3.1.2. Fracture dureza 
Fracture dureza foi medida usando caixa moldada espécimes. A média de altura e largura dos espécimes (N = 10) foram 2,49 e 1,38 milímetros, respectivamente. A profundidade média das sawed faixas foi 1,31 milímetros (S.D. 0,01 milímetros). A carga média de fratura dos espécimes foi 19,38 N (S.D. 2,34 N), que resultou numa média fratura dureza do 5,16 MPa m1 / 2 (S.D. 0,48 MPa m1 / 2; Ver quadro 2). 
Tabela 2. 
Comparação da flexurais fortes do selecionado dentária cerâmicos [13] e [31] 
________________________________________ Flexural força, σ [MPa] 
________________________________________ Weibull força, σ0 [MPa] 
________________________________________ Weibull módulo, metros 
________________________________________ Fracture dureza KIc [MPa m1 / 2] 
________________________________________ 
Zirconia - TZP 91393718 9,42 
Everest HPC 32833917 5,16 
In - Ceram Alumina 6662905 5,00 
IPS Empress 24422899 2,48 
Cerec Mark II 86 88 24 1,29 
IPS Empress 1 84 89 9 1,17 
3.1.3. Chemical solubilidade 
Para os ensaios químicos solubilidade, a 20 disco espécimes foram submetidos aos testes, conforme descrito anteriormente. A média de perda de peso corporal para um único espécime foi 291,5 µ g Isto resultou em um produto químico solubilidade de 7,2 µ g/cm2. 
3,2. Teste de CAD / CAM maquinadas coroas 
Todas as amostras (N = 64) sobreviveram depois de 1.200.000 ciclos de carregamento dinâmico e thermocycling no artificiais orais ambiente. A média de fratura fortes (± S.D.) para as coroas foram: 1789,6 N (658,9), para ZrSiO4, 1715 N (854,3), para Empress 2 coroas e 2415,5 N (801,8), para PFM coroas. A distribuição dos valores podem ser obtidos a partir Fig. 1. A comparação estatística com o teste de Kruskal - Wallis mostrou uma diferença significativa entre a fratura cargas (P = 0,020). A par - sábio comparações múltiplas mostrou diferenças significativas entre a fratura cargas de PFM coroas e Empress 2 coroas (P = 0,032), bem como PFM e ZrSiO4 - coroas (P = 0,007). No entanto, não houve diferença significativa entre ZrSiO4 - coroas e Empress 2 coroas (P = 0,743; Mann - Whitney U - testes). 
  
Display Full Size versão desta imagem (12K) 

Fig. 1. Boxplot da fratura carga dos três tipos de coroas após thermomechanical carga no eixo duplo roer simulador. 
4. Discussão 
Nestes manuscrito dados de ensaios in vitro de um novo encolhimento - livres todos - caco são relatados. 
Uma série de testes estáticos materiais foram realizados para confirmar as propriedades mecânicas deste material sob condições padronizadas. A DIN EN ISO 6872 dobre teste foi realizado em primeiro lugar. A flexurais força de 328 MPa foi demonstrado neste estudo. Embora os valores são muito inferiores aos densa sinterizada cerâmica ou óxido de alumínio óxido de zircónio, este valor coloca o ZrSiO4 cerâmica Evererst HPC nas proximidades de outras cerâmicas, que envolverá um processo sinterização, tais como a IPS Empress 2. Para uma comparação da flexurais dosagens de diferentes materiais cerâmicos utilizados para restaurações dentárias, ver Tabela 2. Esta força faz o material adequado para único dente restaurações de molares e premolars. O elevado Weibull força (339 MPa, ver quadro 2), bem como o elevado Weibull módulo [17] ainda sublinhar a adequação desse material para uso em aplicações única coroa [13] e [31]. 
A fratura obtidos dureza da testados ZrSiO4 cerâmica foi relativamente alta. O valor das ligeiramente acima 5 MPa m1 / 2 foi comparável ao In - Ceram Alumina (ver quadro 2), óxido de um outro baseado caco [13] e [31]. Enquanto nós encontramos um baixo desvio padrão indicando resultados estáveis e uma adequada tamanho amostral ainda podem ser limitações aos nossos dados. Tivemos um entalhe relativamente elevada largura, que pode resultar em grande entalhe raio. A medida fratura dureza é dependente da raiz quadrada do entalhe raio [33], assim, maior radiuses pode resultar em maior dureza fratura valores. A fratura dureza valores como encontrado neste estudo devem ser interpretados com cautela. 
O testado ZrSiO4 cerâmica possui também uma baixa solubilidade química. A solubilidade de 7,2 µ g/cm2 está bem abaixo de um limiar de 10 µ g/cm2 e, assim, ele passa o DIN EN ISO 6872 de referência de 100 µ g/cm2 sem problemas. A química solubilidade desta cerâmica é comparável com óxido de alumínio e de zircónio óxido de ítrio - estabilizado. A baixa solubilidade química é também imprescindível para uma boa biocompatibilidade. A ausência de citotoxicidade, hemotoxicity e genotoxicidade também foram confirmados em testes independentes seguintes DIN EN ISO 10993-1 [12]. 
Fracture força testes de materiais cerâmicos são importantes para a vida útil esperada, com uma baixa probabilidade aceitável de fracasso [35]. A adequação de um novo caco para uma série de indicações podem ser decididas pela comparar resultados obtidos em estudos in vitro com valores conhecidos e aceites por materiais. Portanto, as propriedades físicas do recém - desenvolvidos materiais dentários devem ser testadas antes de poderem ser recomendados para uso clínico. No entanto, as características de insucesso clínico de todos - cerâmicos restaurações não são bem reflectida por parâmetros, como a resistência dos materiais flexurais espécimes ou carga estática - a - fracasso completo de testes restaurações. Estes testes não bem replicar o fracasso mecanismos observados em espécimes clínicos recuperados [20]. Portanto, é mais clinicamente relevantes para testar os espécimes sob fisiológica fadiga carga num simulador roer. Isto permite a avaliação de dentista reparadora sistemas sob condições clínicas relevantes enquanto gravava avaliação tempo [4], [9], [20], [23], [25], [26] e [43]. Foi demonstrado que todos os diferentes - cerâmicos restaurações mostrar fadiga diferentes comportamentos. O roer máquina utilizada neste estudo foi desenvolvido para reproduzir o ambiente no vivo, acrescentando humidade ea uma temperatura controlada para o ensaio condições [23]. Isso é importante quando medição fratura ou fadiga força de dentista cerâmica, como água e temperatura mudança pode atuar em rachadura dicas para diminuir a resistência do vidro e cerâmica [20] e [35]. Segundo a literatura, fadiga produzida por 240.000 ciclos no roer simulador corresponde a um período de 1 ano de serviço clínico [37]. Portanto, 1.200.000 roer ciclos foram realizados para simular um serviço tempo de 5 anos [23] e [26]. Os parâmetros utilizados para a roer máquina foram limitados ao fisiológico valores encontrados na literatura. Embora máxima mordida forças pode facilmente atingir em excesso de 500 N nas áreas posteriores, regularmente masticatory forças podem, mas geralmente não exceder os 49 N carregamento vigor utilizado neste estudo [8], [15] e [38]. 
Os estudos clínicos não são indicados antes a segurança de um reparadora material foi confirmado por testes in vitro. Além disso, são morosos, caros e a padronização do teste parâmetros é difícil. Portanto, mais testes válidos, como a roer simulação efectuada neste estudo são obrigados a incluir outros parâmetros que são suportadas por restaurações na boca. 
Todos coroas sobreviveu a roer simulação, resultando em uma sobrevida de 100% das amostras. Posteriormente, todas as coroas eram carregados a falha em um teste universais máquina. A fratura cargas obtidas por este teste apenas reflectem as cargas às quais fracasso das coroas ocorre. Estes valores variaram entre 966 e 3964 N. Essas cargas não devem ser confundidos com os valores obtidos pelos testes estáticos em materiais espécimes. Crown fracasso cargas fortemente depender da espessura e geometria das coroas. 
A fratura carga do Empress 2 grupo controle mostra que coroas feitas a partir Everest HPC exposição comparável fratura cargas; No entanto, as cargas para o Everest HPC foram obtidos com a cimentação convencional. As altas cargas fratura obtido após carregamento cíclico assim enfatizar que cimentação convencional - como desempenhadas por esses testes - aparece como uma opção viável. Os dados publicados sobre a morder forças indicam que morde a máxima força que pode ocorrer no póstero dentária área varia entre 300 e 880 N [2] e [15]. [18] e [19]. 
Clinical taxas de sobrevivência entre 95 e 97% ao longo de um período de até 7 anos de espera pode ser convencional (expressos e PFM) fixa restaurações [6], [24] e [27] e, após 15-20 anos survivals taxas de 75% Pode - se esperar [7] e [39]. Essas taxas de sobrevivência e de observação vezes tão longe assim ultrapassar as de qualquer sistema tudo - cerâmicos. Clinical sucesso de Empress 2 e In - Ceram todas as restaurações de cerâmica são promissores na sobrevivência das taxas de 78-98% ao longo de 3-6 anos [30], [32] e [40]; No entanto, a verdadeira longo prazo dados ou análises metáforas Como fixos convencionais para restaurações estão ausentes. 
Neste estudo, dentes humanos extraídos foram utilizados como abutments porque seu módulo de elasticidade, vinculação características, condutividade térmica e força são mais estreitas com a situação clínica do que os de metal, plástico e animal dentes. Por outro lado, dentes naturais permitida a cimentação adesiva dos espécimes. Uma tentativa foi feita para selecionar abutments com dimensões comparáveis. Os dentes humanos extraídos foram armazenados em 0,1% timol solução, impedindo - os de secagem para fora e assim se tornando frágil [16] e também inibindo atividade microbiana no armazenamento solução [42]. 
As seguintes conclusões foram tiradas: 
1. O ZrSiO4 cerâmicos testados neste estudo possui propriedades mecânicas que são comparáveis aos dos outros - todos os materiais cerâmicos dentária. 
2. Full coroas feitas de um zircon cerâmica resistir masticatory forças encontrados na área posteriores. 
3. Embora o material parece ser confiável para o uso clínico, para posteriores pleno coroas, avaliação clínica destes todos - cerâmicos restaurações é necessária. 

Referências 
[1] M. A. Awad e J.S. Feine, Measuring doente satisfação com próteses mandibulares, Community Dent Oral Epidemiol 26 (1998), págs. 400-405. Full Text através CrossRef | View Record no Scopus | Citado por em Scopus (35) 
[2] J. F. Bates, G. D. Stafford e A. Harrison, Mastigatórios função de uma revisão da literatura. 2. Velocidade de circulação da mandíbula, taxa de roer e forças desenvolvidas em roer, J Oral Rehabil 2 (1975), págs. 281-301. Full Text através CrossRef | View Record no Scopus | Citado por em Scopus (29) 
[3] A.J. Beaumont Jr., um panorama da estética com parciais removíveis dentadura, Quintessence Int 33 (2002), págs. 747-755. View Record no Scopus | Citado por em Scopus (2) 
[4] M. Behr, M. Rosentritt, A. Leibrock, S. Schneider - Feyrer e G. Handel, In vitro, estudo de fratura força e marginais adaptação de fibra - reforçada adesivos fixados parcial embutimento dentadura, J Dent 27 (1999) , Págs. 163-168. SummaryPlus | Full Text + Links | PDF (830 K) | View Record no Scopus | Citado por em Scopus (37) 
[5] J.R. Binder and W. Die Qual der, Materialvielfalt bei vollkeramischem Zahnersatz, ZWL 3 (2004), pp. 46–51. 
[6] J. Coornaert, P. Adriaens and J De Boever, Long-term clinical study of porcelain-fused-to-gold restorations, J Prosthet Dent 51 (1984), pp. 338–342. Abstract | View Record in Scopus | Cited By in Scopus (38) 
[7] N.H. Creugers, A.F. Kayser and M.A. van’t Hof, A meta-analysis of durability data on conventional fixed bridges, Community Dent Oral Epidemiol 22 (1994), pp. 448–452. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (81) 
[8] J.A. De Boever, W.D. McCall Jr., S. Holden and M.M. Ash Jr., Functional occlusal forces: an investigation by telemetry, J Prosthet Dent 40 (1978), pp. 326–333. Abstract | View Record in Scopus | Cited By in Scopus (54) 
[9] R. DeLong and W.H. Douglas, Development of an artificial oral environment for the testing of dental restoratives: bi-axial force and movement control, J Dent Res 62 (1983), pp. 32–36. View Record in Scopus | Cited By in Scopus (93) 
[10] W.J. Dunn, D.F. Murchison and J.C. Broome, Esthetics: patients’ perceptions of dental attractiveness, J Prosthodont 5 (1996), pp. 166–171. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (22) 
[11] R.L. Engelmeier, Complete-denture esthetics, Dent Clin North Am 40 (1996), pp. 71–84. View Record in Scopus | Cited By in Scopus (7) 
[12] Fiedler S, Ohmacht M. Zertifikat Biokompatibilität des Prüfmaterials KaVo Everest HPC. Ochsenhausen: MDT Medical Device Testing; 2005. 
[13] H. Fischer and R. Marx, Festigkeit von Dentalkeramik, ZWR 109 (2000), pp. 240–244. 
[14] R. Gerhard and N. Sawyer, Dentures to harmonize with heavily pigmented tissues, J Am Dent Assoc 73 (1966), pp. 94–95. View Record in Scopus | Cited By in Scopus (3) 
[15] C.H. Gibbs, P.E. Mahan, A. Mauderli, H.C. Lundeen and E.K. Walsh, Limits of human bite strength, J Prosthet Dent 56 (1986), pp. 226–229. Abstract | View Record in Scopus | Cited By in Scopus (47) 
[16] A.R. Helfer, S. Melnick and H. Schilder, Determination of the moisture content of vital and pulpless teeth, Oral Surg Oral Med Oral Pathol 34 (1972), pp. 661–670. Abstract | View Record in Scopus | Cited By in Scopus (51) 
[17] V.D. Hennige, J. Hausselt, H.J. Ritzhaupt-Kleissl and T. Windmann, Shrinkage-free ZrSiO4-ceramics: characterisation and applications, J Eur Ceram Soc 19 (1999), pp. 2901–2908. SummaryPlus | Full Text + Links | PDF (270 K) | View Record in Scopus | Cited By in Scopus (9) 
[18] H.F. Kappert, H. Knode and L. Manzotti, Metal-free bridges in posterior region, Dent Labor (Munch)38 (1990), pp. 177–183. View Record in Scopus | Cited By in Scopus (8) 
[19] H.F. Kappert, H. Knode and R. Schultheiss, Strength of the In-Ceram system under mechanical loads and thermocycling in artificial saliva, Dtsch Zahnarztl Z 46 (1991), pp. 129–131. View Record in Scopus | Cited By in Scopus (21) 
[20] J.R. Kelly, Clinically relevant approach to failure testing of all-ceramic restorations, J Prosthet Dent 81 (1999), pp. 652–661. SummaryPlus | Full Text + Links | PDF (296 K) | View Record in Scopus | Cited By in Scopus (116) 
[21] J.R. Kelly, J.A. Tesk and J.A. Sorensen, Failure of all-ceramic fixed partial dentures in vitro and in vivo: analysis and modeling, J Dent Res 74 (1995), pp. 1253–1258. View Record in Scopus | Cited By in Scopus (93) 
[22] M. Kern, W.H. Douglas, T. Fechtig, J.R. Strub and R. DeLong, Fracture strength of all-porcelain, resin-bonded bridges after testing in an artificial oral environment, J Dent 21 (1993), pp. 117–121. Abstract | View Record in Scopus | Cited By in Scopus (26) 
[23] M. Kern, J.R. Strub and X.Y. Lu, Wear of composite resin veneering materials in a dual-axis chewing simulator, J Oral Rehabil 26 (1999), pp. 372–378. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (49) 
[24] T. Kerschbaum and R. Voss, Comparison of cast and metal–ceramic veneer crowns: results of a follow-up study in wearers of partial prostheses, Dtsch Zahnarztl Z 32 (1977), pp. 200–206. View Record in Scopus | Cited By in Scopus (4) 
[25] S.O. Koutayas, M. Kern, F. Ferraresso and J.R. Stub, Influence of design and mode of loading on the fracture strength of all- ceramic resin-bonded fixed partial dentures: an in vitro study in a dual-axis chewing simulator, J Prosthet Dent 83 (2000), pp. 540–547. SummaryPlus | Full Text + Links | PDF (265 K) | View Record in Scopus | Cited By in Scopus (15) 
[26] I. Krejci, T. Reich, F. Lutz and M. Albertoni, An in vitro test procedure for evaluating dental restoration systems. 1. A computer-controlled mastication simulator, Schweiz Monatsschr Zahnmed 100 (1990), pp. 953–960. View Record in Scopus | Cited By in Scopus (42) 
[27] P.J. Leempoel, S. Eschen, A.F. De Haan and M.A. Van’t Hof, An evaluation of crowns and bridges in a general dental practice, J Oral Rehabil 12 (1985), pp . 515–528. Full Text via CrossRef | View Record in Scopus | Cited By in Scopus (42) 
[28] R.G. Luthardt, M. Holzhuter and O. Sandkuhl et al., Reliability and properties of ground Y-TZP-zirconia ceramics, J Dent Res 81 (2002), pp. 487–491. View Record in Scopus | Cited By in Scopus (24) 
[29] R.G. Luthardt, M.S. Holzhuter, H. Rudolph, V. Herold and M.H. Walter, CAD/CAM-machining effects on Y-TZP zirconia, Dent Mater 20 (2004), pp. 655–662. SummaryPlus | Full Text + Links | PDF (433 K) | View Record in Scopus | Cited By in Scopus (16) 
[30] P. Marquardt and J.R. Strub, Survival rates of IPS Empress 2 all-ceramic crowns and fixed partial dentures: results of a 5-year prospective clinical study, Quintessence Int 37 (2006), pp. 253–259. View Record in Scopus | Cited By in Scopus (11) 
[31] R. Marx, H. Fischer, M. Weber and F. Jungwirth, Crack parameters and Weibull moduli: subcritical crack growth and long-time durability of all-ceramic materials, Dtsch Zahnärztl Z 56 (2001), pp. 90–98. 
[32] E.A. McLaren and S.N. White, Survival of In-Ceram crowns in a private practice: a prospective clinical trial, J Prosthet Dent 83 (2000), pp. 216–222. Abstract | PDF (692 K) | View Record in Scopus | Cited By in Scopus (40) 
[33] D. Munz and T. Fett, Fracture mechanics. In: D. Munz and T. Fett, Editors, Ceramics: mechanical properties, failure behaviour, materials selection, Springer, Berlin, Heidelberg (1999), pp. 19–51. 
[34] W. Riedling and H.F. Kappert, Longevity of Hi-Ceram full ceramic bridges, Quintessenz 38 (1987), pp. 1493–1498. View Record in Scopus | Cited By in Scopus (1) 
[35] J.E. Ritter, Predicting lifetimes of materials and material structures, Dent Mater 11 (1995), pp. 142–146. SummaryPlus | Full Text + Links | PDF (617 K) | View Record in Scopus | Cited By in Scopus (25) 
[36] A.S. Rizkalla and D.W. Jones, Mechanical properties of commercial high strength ceramic core materials, Dent Mater 20 (2004), pp. 207–212. SummaryPlus | Full Text + Links | PDF (338 K) | View Record in Scopus | Cited By in Scopus (10) 
[37] R.L. Sakaguchi, W.H. Douglas, R. DeLong and M.R. Pintado, The wear of a posterior composite in an artificial mouth: a clinical correlation, Dent Mater 2 (1986), pp. 235–240. Abstract | View Record in Scopus | Cited By in Scopus (44) 
[38] H. Schwickerath and I. Coca, Single crowns of glass–ceramic, Phillip J Restaur Zahnmed 4 (1987), pp. 336–338. View Record in Scopus | Cited By in Scopus (4) 
[39] M.S. Scurria, J.D. Bader and D.A. Shugars, Meta-analysis of fixed partial denture survival: prostheses and abutments, J Prosthet Dent 79 (1998), pp. 459–464. SummaryPlus | Full Text + Links | PDF (158 K) | View Record in Scopus | Cited By in Scopus (74) 
[40] B.S. Segal, Retrospective assessment of 546 all-ceramic anterior and posterior crowns in a general practice, J Prosthet Dent 85 (2001), pp. 544–550. Abstract | PDF (193 K) | View Record in Scopus | Cited By in Scopus (24) 
[41] P.N. Sellen, D.C. Jagger and A. Harrison, Methods used to select artificial anterior teeth for the edentulous patient: a historical overview, Int J Prosthodont 12 (1999), pp. 51–58. View Record in Scopus | Cited By in Scopus (14) 
[42] O. Sparrius and E.S. Grossman, Marginal leakage of composite resin restorations in combination with dentinal and enamel bonding agents, J Prosthet Dent 61 (1989), pp. 678–684. Abstract | View Record in Scopus | Cited By in Scopus (10) 
[43] J.R. Strub and S.M. Beschnidt, Fracture strength of 5 different all-ceramic crown systems, Int J Prosthodont 11 (1998), pp. 602–609. View Record in Scopus | Cited By in Scopus (29) 


  Corresponding author at: Department of Prosthodontics, University Hospital, Hugstetter Sr. 55, D-79106 Freiburg, Germany. Tel.: +49 761 270 4906; fax: +49 761 270 4824. 



Dental Materials 
Volume 23, Issue 7, July 2007, Pages 785-791